# Appendix A: Measurement Quality Objective Tables

# Table A2: Measurement Quality Objectives* – Conventional Analytes in Fresh and Marine Waters – Total, Fixed, Suspended and Volatile Solids

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Laboratory Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% (n/a if native concentration of either sample<RL) |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | RPD<25% (n/a if native concentration of either sample<RL) |
| Field Blank, Equipment Blank | Per method | <RL for target analyte |

*Unless method specifies more stringent requirements

Table A4: Measurement Quality Objectives* - Conventional Analytes in Fresh and Marine Sediments

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 10 analytical runs (as applicable) | 80-120% recovery |
| Laboratory Blank | TOC only: one per analytical batch (n/a for others) | <RL or <30% of lowest sample |
| Reference Material | TOC only: one per 20 samples or per analytical batch, whichever is more frequent (n/a for others) | 80-120% recovery |
| Matrix Spike | n/a | n/a |
| Matrix Spike Duplicate | n/a | n/a |
| Laboratory Duplicate | One per analytical batch | RPD<25% (n/a if native concentration of either sample<RL) |
| Surrogate or Internal Standard | n/a | n/a |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | RPD<25% (n/a if native concentration of either sample<RL) |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL or <30% of lowest sample |

*Unless method specifies more stringent requirements

# Table A5: Measurement Quality Objectives* – Inorganic Analytes in Fresh and Marine Sediments and Tissues

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 10 analytical runs | 80-120% recovery |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Reference Material | Per 20 samples or per analytical batch, whichever is more frequent | 75-125% recovery (70-130% for MMHg) |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 75-125% recovery (70-130% for MMHg) |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | 75-125% recovery (70-130% for MMHg); RPD<25% |
| Laboratory Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% (n/a if native concentration of either sample<RL) |
| Internal Standard | Accompanying every analytical run when method appropriate | 60-125% recovery |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | RPD<25% (n/a if native concentration of either sample<RL), unless otherwise specified by method |
| Field Blank, Equipment Blank | Per method | Blanks<RL for target analyte |

*Unless method specifies more stringent requirements

# Table A6: Measurement Quality Objectives* – Volatile Organic Compounds in Fresh and Marine Sediments

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 12 hours | RF for SPCCs same as initial calibration; RF of CCVs must be within 20% of initial calibration |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Reference Material | Method Validation: as many as required to assess accuracy and precision of method before routine analysis of samples; Routine Accuracy Assessment: per 20 samples or per analytical batch (preferably blind) | 70-130% recovery if certified; otherwise 50-150% recovery |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 50-150% recovery, or based on 3x the standard deviation of laboratory's actual method recoveries |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% |
| Laboratory Duplicate | Per method | Per method |
| Surrogate or Internal Standard | Per method | Per method |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | Per method |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL for target analyte |

*Unless method specifies more stringent requirements

# Table A7: Measurement Quality Objectives* – Semi-Volatile Organic Compounds in Fresh and Marine Sediments

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 12 h | RF for SPCCs same as initial calibration; RF of CCVs must be within 20% of initial calibration |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Reference Material | Method Validation: as many as required to assess accuracy and precision of method before routine analysis of samples; Routine Accuracy Assessment: per 20 samples or per analytical batch (preferably blind) | 70-130% recovery if certified; otherwise, 50-150% recovery |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 50-150% recovery, or based on 3x the standard deviation of laboratory's actual method recoveries |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% |
| Laboratory Duplicate | Per method | Per method |
| Surrogate or Internal Standard | Per method | Per method |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | Per method |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL for target analyte |

*Unless method specifies more stringent requirements

# Table A8: Measurement Quality Objectives* – Synthetic Organic Compounds in Fresh and Marine Sediments and Tissues

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 10 analytical runs | Water: 85-115% recovery Sediment: 85-115% recovery Tissue: 75-125% |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analytes |
| Reference Material | Method Validation: as many as required to assess accuracy and precision of method before routine analysis of samples; Routine Accuracy Assessment: per 20 samples or per analytical batch (preferably blind) | 70-130% recovery if certified; otherwise, 50-150% recovery |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 50-150% recovery, or based on 3x the standard deviation of laboratory's actual method recoveries |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% |
| Laboratory Duplicate | Per method | Water: RPD<25% (n/a if native concentration of either sample<RL) Sediment: Per method Tissue: Per method |
| Surrogate or Internal Standard | Per method | Per method |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | Per method |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL for target analytes |

* Unless method specifies more stringent requirements. ELISA results must be assessed against kit requirements