Appendix II

QUALITY CONTROL DATA

FOR ANALYTICAL CHEMISTRY 

OF FISH TISSUE SAMPLES

Table A1-1. SWAMP Quality Control (QC) Requirements* for PCB Congeners in Tissue.

	Type of QC Sample
	Purpose
	Required Frequency
	Control Limit

	Method Blank
	Reagent contamination
	1 per 20 samples or 1 per batch, whichever is more frequent
	< MDL for analyte of interest

	Calibration standards
	Establish relationship between instrument response and target analyte concentration
	As specified in method or laboratory SOP
	Linear regression r>0.995

	Calibration verification standard
	Assess instrument drift
	Every 10 samples
	85-115% recovery

	Certified reference material
	Accuracy
	1 per 20 samples or 1 per batch, whichever is more frequent
	70-130% recovery of the 95% confidence level

	Matrix spike (MS) and matrix spike duplicate (MSD)
	Matrix effects and method performance
	1 MS and 1 MSD 20 samples or 1 MS/MSD pair per batch, whichever is more frequent
	50-150% recovery and RPD<25%

	Surrogate spikes
	Assess method performance and estimate recovery of target analytes
	Added to every calibration standard, sample, and blank
	As specified in method or by project manager

	Field replicate
	Method precision, homogeneity of sample
	1 per 20 samples or 1 per batch, whichever is more frequent
	RPD<25% if samples is 10 times MDL


Table A1-2. SWAMP Quality Control Requirements* for Total Mercury in Tissue.

	Type of QC Sample
	Purpose
	Required Frequency
	Control Limit

	Method Blank
	Reagent contamination
	1 per 20 samples or 1 per batch, whichever is more frequent
	< MDL for analyte of interest

	Calibration standards
	Establish relationship between instrument response and target analyte concentration
	As specified in method or laboratory SOP
	Linear regression r>0.995

	Calibration verification standard
	Assess instrument drift
	Every 10 samples
	80-120% recovery

	Certified reference material
	Accuracy
	1 per 20 samples or 1 per batch, whichever is more frequent
	75-125% recovery

	Matrix spike (MS) and matrix spike duplicate (MSD)
	Matrix effects and method performance
	1 MS and 1 MSD 20 samples or 1 MS/MSD pair per batch, whichever is more frequent
	75-125% recovery and RPD<25%

	Field replicate
	Method precision, homogeneity of sample
	1 per 20 samples or 1 per batch, whichever is more frequent
	RPD<25% if samples is 10 times MDL


*The two tables represent some of the QC requirements for tissue analysis under the SWAMP QA program

RPD = relative percent difference (sample difference divided by sample mean and multiplied by 100%)

SOP = standard operating procedure

MDL = method detection limit

RL = reporting limit
Table A1-3. Method blank results for PCB Congeners. Results reported in ppb (wet weight). MDL = 0.1 ppb and RL = 0.2 ppb.

	 QA-Designated
	 
	Method Blank
	Method Blank

	Batch Identification
	PCB Congener
	Sample Result
	Sample Result

	TSM2000
	18
	ND
	ND

	TSM2000
	28
	0.096
	0.088

	TSM2000
	31
	0.099
	0.091

	TSM2000
	99
	0.044
	0.060

	TSM2000
	118
	0.212
	0.237

	TSM2000
	128
	ND
	ND

	TSM2000
	138
	0.112
	0.129

	TSM2000
	149
	ND
	0.044

	TSM2000
	153
	0.053
	0.065

	TSM2000
	180
	ND
	ND

	TSM2000
	194
	ND
	ND

	TSM2000
	195
	ND
	ND

	TSM2000
	201
	ND
	ND

	TSM2000
	203
	ND
	ND

	
	
	
	

	TSM2001
	18
	ND
	NA

	TSM2001
	28
	ND
	NA

	TSM2001
	31
	ND
	NA

	TSM2001
	99
	ND
	NA

	TSM2001
	118
	0.296
	NA

	TSM2001
	128
	ND
	NA

	TSM2001
	138
	0.198
	NA

	TSM2001
	149
	ND
	NA

	TSM2001
	153
	ND
	NA

	TSM2001
	180
	ND
	NA

	TSM2001
	194
	ND
	NA

	TSM2001
	195
	ND
	NA

	TSM2001
	201
	ND
	NA

	TSM2001
	203
	ND
	NA


NA = not analyzed

ND = non-detect (below the MDL)
Table A1-4. Method blank results for Total Mercury. Results reported in ppm (dry weight).  MDL = 0.011 and RL = 0.024 ppm.(dry weight).

	Lab Batch 

Identification
	Lab Sample ID
	Result

	TSM00THg1
	MB-1
	ND

	TSM00THg1
	MB-2
	ND

	TSM00THg1
	MB-3
	ND

	TSM00THg2
	MB-1
	ND

	TSM00THg2
	MB-2
	ND

	TSM00THg2
	MB-3
	ND

	TSM00THg3
	MB-1
	ND

	TSM00THg3
	MB-2
	ND

	TSM00THg3
	MB-3
	ND

	TSM01THg1
	MB-1
	ND

	TSM01THg1
	MB-2
	ND

	TSM01THg1
	MB-3
	ND

	TSM01THg2
	MB-1
	ND

	TSM01THg2
	MB-2
	ND

	TSM01THg2
	MB-3
	ND


ND = non-detect (below the MDL)

Table A1-5. Certified reference material results for PCB Congeners. NIST SRM-2978 mussel tissue. Results reported in ppb (dry weight).

	QA-Designated
	 
	 
	 
	95% CI
	 
	 
	 
	 
	 

	Batch
	PCB 
	Certified
	
	Ranges
	70-130% of the 95%
	
	
	 

	Identification
	Congener  
	Value
	+/-
	Lower
	Upper
	Confidence Interval
	Result
	Result
	RPD

	TSM2000
	28
	7.91
	0.9
	7.01
	8.81
	4.91
	11.45
	7.45
	7.78
	4.3

	TSM2000
	31
	21.4
	0.43
	21.0
	21.83
	14.68
	28.38
	6.05
	6.34
	4.7

	TSM2000
	44
	11.8
	0.64
	11.16
	12.44
	7.81
	16.17
	11.3
	10.9
	3.9

	TSM2000
	49
	16.84
	0.86
	16.0
	17.7
	11.19
	23.01
	14.6
	13.2
	10.0

	TSM2000
	52
	17.7
	2.8
	14.9
	20.5
	10.43
	26.65
	17.8
	16.3
	8.7

	TSM2000
	66
	18.4
	1.5
	16.9
	19.9
	11.83
	25.87
	24.5
	22.5
	8.8

	TSM2000
	87
	10.2
	0.29
	9.91
	10.49
	6.94
	13.64
	12.3
	11.9
	3.2

	TSM2000
	95
	20.8
	2.1
	18.7
	22.9
	13.09
	29.77
	23.0
	19.4
	17.3

	TSM2000
	99
	18.84
	0.44
	18.4
	19.28
	12.88
	25.06
	20.2
	19.0
	6.0

	TSM2000
	101
	35.9
	1.6
	34.3
	37.5
	24.01
	48.75
	46.6
	45.2
	3.1

	TSM2000
	105
	10.8
	0.45
	10.35
	11.25
	7.25
	14.63
	7.28
	12.8
	54.7

	TSM2000
	110
	35.34
	0.71
	34.63
	36.05
	24.24
	46.87
	37.0
	42.3
	13.3

	TSM2000
	118
	35.1
	1.0
	34.1
	36.1
	23.87
	46.93
	40.0
	41.7
	4.1

	TSM2000
	128
	5.25
	0.17
	5.08
	5.42
	3.56
	7.05
	4.90
	5.47
	10.9

	TSM2000
	138
	35.7
	1.5
	34.2
	37.2
	23.94
	48.36
	46.5
	46.5
	0.1

	TSM2000
	149
	34.73
	0.69
	34.0
	35.42
	23.83
	46.05
	29.8
	29.0
	2.6

	TSM2000
	151
	10.92
	0.25
	10.67
	11.17
	7.47
	14.52
	11.1
	11.3
	1.8

	TSM2000
	153
	56.9
	3.5
	53.4
	60.4
	37.38
	78.52
	55.3
	54.7
	1.1

	TSM2000
	156
	1.97
	0.11
	1.86
	2.08
	1.30
	2.70
	1.99
	2.09
	4.7

	TSM2000
	180
	7.81
	0.63
	7.18
	8.44
	5.03
	10.97
	6.48
	6.17
	4.9

	TSM2000
	183
	5.25
	0.15
	5.10
	5.40
	3.57
	7.02
	4.65
	4.14
	11.6

	TSM2000
	187
	16.7
	1.3
	15.40
	18.00
	10.78
	23.40
	16.9
	16.4
	3.0


Table A1-5 (continued). Certified reference material results for PCB Congeners. NIST SRM-2978 mussel tissue. Results reported in ppb (dry weight).

	QA-Designated
	 
	 
	 
	95% CI
	 
	 
	 

	Batch
	PCB 
	Certified
	
	Ranges
	70-130% of the 95%
	

	Identification
	Congener  
	Value
	+/-
	Lower
	Upper
	Confidence Interval
	Result

	TSM2001
	28
	7.91
	0.9
	7.01
	8.81
	4.91
	11.45
	5.30

	TSM2001
	31
	21.4
	0.43
	21.0
	21.83
	14.68
	28.38
	5.26

	TSM2001
	44
	11.8
	0.64
	11.16
	12.44
	7.81
	16.17
	11.0

	TSM2001
	49
	16.84
	0.86
	16.0
	17.7
	11.19
	23.01
	12.4

	TSM2001
	52
	17.7
	2.8
	14.9
	20.5
	10.43
	26.65
	15.5

	TSM2001
	66
	18.4
	1.5
	16.9
	19.9
	11.83
	25.87
	17.6

	TSM2001
	87
	10.2
	0.29
	9.91
	10.49
	6.94
	13.64
	10.4

	TSM2001
	95
	20.8
	2.1
	18.7
	22.9
	13.09
	29.77
	19.6

	TSM2001
	99
	18.84
	0.44
	18.4
	19.28
	12.88
	25.06
	18.3

	TSM2001
	101
	35.9
	1.6
	34.3
	37.5
	24.01
	48.75
	42.3

	TSM2001
	105
	10.8
	0.45
	10.35
	11.25
	7.25
	14.63
	9.19

	TSM2001
	110
	35.34
	0.71
	34.63
	36.05
	24.24
	46.87
	37.4

	TSM2001
	118
	35.1
	1.0
	34.1
	36.1
	23.87
	46.93
	36

	TSM2001
	128
	5.25
	0.17
	5.08
	5.42
	3.56
	7.05
	4.86

	TSM2001
	138
	35.7
	1.5
	34.2
	37.2
	23.94
	48.36
	46.6

	TSM2001
	149
	34.73
	0.69
	34.0
	35.42
	23.83
	46.05
	29.9

	TSM2001
	151
	10.92
	0.25
	10.67
	11.17
	7.47
	14.52
	9.54

	TSM2001
	153
	56.9
	3.5
	53.4
	60.4
	37.38
	78.52
	59.8

	TSM2001
	156
	1.97
	0.11
	1.86
	2.08
	1.30
	2.70
	1.76

	TSM2001
	180
	7.81
	0.63
	7.18
	8.44
	5.03
	10.97
	5.64

	TSM2001
	183
	5.25
	0.15
	5.10
	5.40
	3.57
	7.02
	4.43

	TSM2001
	187
	16.7
	1.3
	15.40
	18.00
	10.78
	23.40
	16.1


Table A1-6. Certified reference material results for Total Mercury. NIST DORM-2 dogfish muscle certified value = 4.64 ± 0.24 ppm. Results reported in ppm (dry weight).

	Lab 

Batch Identification
	Result
	Percent Recovery

	TSM00THg1
	4.71
	101.5

	TSM00THg2
	4.61
	99.4

	TSM00THg3
	4.46
	96.1

	TSM01THg1
	5.12
	110.3

	TSM01THg2
	4.85
	104.5


Figure A1-1. Control chart for percent recovery in DORM-2 for Total Mercury.


[image: image1.wmf]Certified Reference Material Total Mercury

NIST DORM-2 Dogfish Muscle certified value = 4.64 +/- 0.24 ppm

control limit = 75-125% recovery

n = 5, sample mean = 4.75 (102.4%)

70

80

90

100

110

120

130

Batch Identification

Percent Recovery


Table A1-7. Replicate analysis samples for PCB Congeners. RPD = relative percent difference (sample difference divided by sample mean and multiplied by 100%). Results reported in ppb (wet weight).
	QA-Designated
	Parent 
	 
	 Parent
	 Replicate
	 
	 

	Batch
	Sample
	 
	Sample
	Sample
	
	Sample

	Identification
	Identification
	PCB Congener 
	Result
	Result
	RPD
	<10xMDL

	TSM2000
	384.010.F.00
	18
	0.314
	0.334
	6.2
	x

	TSM2000
	384.010.F.00
	28
	1.17
	1.18
	0.2
	 

	TSM2000
	384.010.F.00
	31
	1.01
	1.01
	0.2
	 

	TSM2000
	384.010.F.00
	99
	6.26
	6.16
	1.6
	 

	TSM2000
	384.010.F.00
	118
	15.3
	15.9
	3.6
	 

	TSM2000
	384.010.F.00
	128
	3.89
	4.95
	24.0
	 

	TSM2000
	384.010.F.00
	138
	37.8
	39.4
	4.0
	 

	TSM2000
	384.010.F.00
	149
	11.2
	11.3
	1.4
	 

	TSM2000
	384.010.F.00
	153
	44.0
	46.5
	5.6
	 

	TSM2000
	384.010.F.00
	180
	21.6
	22.5
	4.1
	 

	TSM2000
	384.010.F.00
	194
	3.09
	3.09
	0.1
	 

	TSM2000
	384.010.F.00
	195
	0.596
	0.497
	18.1
	x

	TSM2000
	384.010.F.00
	201
	4.85
	4.97
	2.5
	 

	TSM2000
	384.010.F.00
	203
	3.51
	3.57
	1.6
	 

	 
	 
	 
	 
	 
	 
	 

	TSM2000
	384.001.F.00
	18
	0.135
	0.126
	7.2
	x

	TSM2000
	384.001.F.00
	28
	0.288
	0.302
	5.0
	x

	TSM2000
	384.001.F.00
	31
	0.230
	0.277
	18.9
	x

	TSM2000
	384.001.F.00
	99
	0.479
	0.509
	6.0
	x

	TSM2000
	384.001.F.00
	118
	1.04
	1.14
	9.2
	 

	TSM2000
	384.001.F.00
	128
	0.240
	0.264
	9.7
	x

	TSM2000
	384.001.F.00
	138
	1.66
	1.81
	8.4
	 


Table A1-7 (continued). Replicate analysis samples for PCB Congeners. RPD = relative percent difference (sample difference divided by sample mean and multiplied by 100%). Results reported in ppb (wet weight).
	QA-Designated
	Parent 
	 
	 Parent
	Replicate
	 
	 

	Batch
	Sample
	 
	Sample
	Sample
	
	Sample

	Identification
	Identification
	PCB Congener 
	Result
	Result
	RPD
	<10xMDL

	TSM2000
	384.001.F.00
	149
	0.752
	0.818
	8.4
	x

	TSM2000
	384.001.F.00
	153
	1.51
	1.63
	7.5
	 

	TSM2000
	384.001.F.00
	180
	0.662
	0.696
	4.9
	x

	TSM2000
	384.001.F.00
	194
	0.107
	0.105
	2.3
	x

	TSM2000
	384.001.F.00
	195
	ND
	ND
	ND
	 

	TSM2000
	384.001.F.00
	201
	0.174
	0.173
	0.8
	x

	TSM2000
	384.001.F.00
	203
	0.122
	0.169
	32.6
	x

	 
	 
	 
	 
	 
	 
	 

	TSM2001
	013.002.F.01  
	18
	ND
	ND
	ND
	x

	TSM2001
	013.002.F.01  
	28
	0.102
	0.276
	92.1
	x

	TSM2001
	013.002.F.01  
	31
	0.093
	0.224
	82.6
	x

	TSM2001
	013.002.F.01  
	99
	0.235
	0.273
	15.0
	x

	TSM2001
	013.002.F.01  
	118
	0.655
	0.834
	24.0
	x

	TSM2001
	013.002.F.01  
	128
	0.113
	0.112
	0.9
	x

	TSM2001
	013.002.F.01  
	138
	0.950
	0.846
	11.6
	x

	TSM2001
	013.002.F.01  
	149
	0.298
	0.287
	3.8
	x

	TSM2001
	013.002.F.01  
	153
	0.750
	0.680
	9.8
	x

	TSM2001
	013.002.F.01  
	180
	0.251
	0.238
	5.3
	x

	TSM2001
	013.002.F.01  
	194
	0.038
	0.024
	<MDL
	x

	TSM2001
	013.002.F.01  
	195
	ND
	ND
	ND
	x

	TSM2001
	013.002.F.01  
	201
	0.031
	0.027
	<MDL
	x

	TSM2001
	013.002.F.01  
	203
	0.043
	0.039
	<MDL
	x


Figure A1-2. Control chart for replicate sample analysis RPD for PCB Congeners. Only results >10 times the MDL are shown on control chart.


[image: image2.wmf]Replicate Sample Analysis PCB Congeners

control limit = 25%, if sample result >10x MDL

n = 15, mean RPD = 4.9

-2.0

5.0

12.0

19.0

26.0

33.0

PCB

18

PCB

28

PCB

31

PCB

99

PCB

118

PCB

128

PCB

138

PCB

149

PCB

153

PCB

180

PCB

194

PCB

195

PCB

201

PCB

203

PCB Congener

Relative Percent 

Difference


Table A1-8. Replicate analysis samples for Total Mercury. RPD = relative percent difference (sample difference divided by sample mean and multiplied by 100%). Results reported in ppm (dry weight).
	Lab Batch 

Identification
	Parent Sample 

Identification
	Parent Sample 

Result
	Replicate Sample 

Result
	RPD

	TSM00THg1
	049.002.F.00
	5.26
	5.53
	5

	TSM00THg2
	107.006.F.00
	3.68
	3.63
	1.37

	TSM00THg3
	384.006.F.00
	0.914
	0.861
	5.97

	TSM01THg1
	064.001.F.01
	0.449
	0.456
	1.55

	TSM01THg2
	387.009.F.01
	1.14
	1.11
	2.67


Figure A1-3. Control chart for replicate sample analysis RPD for Total Mercury.


[image: image3.wmf]Replicate Sample Analysis Total Mercury

control limit = 25% RPD, if sample result >10x MDL

n = 10, mean RPD = 3.3

0

5

10

15

20

25

30

TSM00THg1

TSM00THg2

TSM00THg3

TSM01THg1

TSM01THg2

Batch Identification

Relative Percent Difference


Table A1-10. Matrix spike (MS) and matrix spike duplicate (MSD) analysis samples for PCB Congeners. Results reported in ppb (wet weight). **Spike-level too low, result discarded.
	QA-Designated
	Parent
	 
	 
	 
	 
	 

	Batch
	Sample
	
	MS
	Percent
	MSD
	Percent

	Identification
	Identification
	PCB Congener
	Result
	Recovery
	Result
	Recovery

	TSM2000
	384.012.F.00
	18
	2.06
	91.3
	2.00
	87.9

	TSM2000
	384.012.F.00
	28
	3.23
	121
	3.18
	118

	TSM2000
	384.012.F.00
	31
	2.86
	108
	2.50
	88.8

	TSM2000
	384.012.F.00
	99
	6.23
	110
	6.16
	106

	TSM2000
	384.012.F.00
	118
	12.4
	**
	12.5
	**

	TSM2000
	384.012.F.00
	128
	4.96
	117
	5.16
	126

	TSM2000
	384.012.F.00
	138
	27.1
	**
	26.7
	**

	TSM2000
	384.012.F.00
	149
	8.29
	115
	8.26
	113

	TSM2000
	384.012.F.00
	153
	31.2
	**
	30.7
	**

	TSM2000
	384.012.F.00
	180
	17.1
	**
	17.0
	**

	TSM2000
	384.012.F.00
	194
	4.23
	108
	4.32
	112

	TSM2000
	384.012.F.00
	195
	2.98
	110
	2.97
	109

	TSM2000
	384.012.F.00
	201
	5.38
	111
	5.36
	109

	TSM2000
	384.012.F.00
	203
	3.46
	57.0
	4.14
	90.3

	 
	 
	 
	 
	 
	 
	 

	TSM2000
	384.002.F.00
	18
	1.90
	89.6
	2.00
	95.7

	TSM2000
	384.002.F.00
	28
	2.59
	105
	2.54
	103

	TSM2000
	384.002.F.00
	31
	2.10
	91.3
	2.09
	91.7

	TSM2000
	384.002.F.00
	99
	3.65
	97.1
	3.57
	94.0

	TSM2000
	384.002.F.00
	118
	6.41
	121
	6.29
	116

	TSM2000
	384.002.F.00
	128
	2.97
	105
	2.93
	104

	TSM2000
	384.002.F.00
	138
	8.38
	112
	8.20
	104

	TSM2000
	384.002.F.00
	149
	3.32
	97.3
	3.24
	94.2

	TSM2000
	384.002.F.00
	153
	10.6
	126
	10.3
	112

	TSM2000
	384.002.F.00
	180
	5.35
	112
	5.25
	108

	TSM2000
	384.002.F.00
	194
	2.31
	90.4
	2.32
	92.0

	TSM2000
	384.002.F.00
	195
	2.59
	121
	2.59
	122

	TSM2000
	384.002.F.00
	201
	2.71
	100
	2.69
	100

	TSM2000
	384.002.F.00
	203
	2.40
	89.1
	2.18
	79.0

	 
	 
	 
	 
	 
	 
	 

	TSM2001
	064.8.F.01
	18
	1.61
	64.7
	1.66
	67.4

	TSM2001
	064.8.F.01
	28
	1.95
	95.3
	1.95
	95.4

	TSM2001
	064.8.F.01
	31
	2.16
	103
	2.18
	104

	TSM2001
	064.8.F.01
	99
	2.60
	98.5
	2.49
	93.3

	TSM2001
	064.8.F.01
	118
	3.42
	105
	3.26
	96.9

	TSM2001
	064.8.F.01
	128
	1.92
	72.0
	1.92
	72.1

	TSM2001
	064.8.F.01
	138
	6.87
	125
	6.39
	101

	TSM2001
	064.8.F.01
	149
	4.65
	122
	4.29
	104

	TSM2001
	064.8.F.01
	153
	7.47
	137
	6.87
	107

	TSM2001
	064.8.F.01
	180
	5.28
	117
	4.84
	95.2

	TSM2001
	064.8.F.01
	194
	2.25
	90.3
	2.14
	85.2

	TSM2001
	064.8.F.01
	195
	2.03
	93.3
	1.98
	90.6

	TSM2001
	064.8.F.01
	201
	2.55
	97.2
	2.40
	90.0

	TSM2001
	064.8.F.01
	203
	2.65
	106
	2.48
	97.6


Figure A1-4. Control chart for MS/MSD sample analysis for PCB.


[image: image4.wmf]Matrix Spike Sample Analysis PCB Congeners

control limit = 50-150% recovery

n = 76

25

50

75

100

125

150

175

PCB

18

PCB

28

PCB

31

PCB

99

PCB

118

PCB

128

PCB

138

PCB

149

PCB

153

PCB

180

PCB

194

PCB

195

PCB

201

PCB

203

PCB Congener

Percent Recovery


Table A1-11. Matrix spike (MS) and matrix spike duplicate (MSD) analysis samples for Total Mercury. Results reported in ppm (dry weight).
	Lab Batch

Identification
	Parent Sample 

Identification
	Parent Sample 

Result
	MS

Result
	Percent 

Recovery
	MSD 

Result
	Percent

Recovery

	TSM00THg1
	049.002.F.00
	5.26
	21.8
	101.0
	22.2
	103.0

	TSM00THg2
	107.006.F.00
	3.68
	14.0
	93.4
	14.7
	100.0

	TSM00THg3
	384.006.F.00
	0.91
	3.49
	96.5
	3.54
	98.3

	TSM01THg1
	064.001.F.01
	0.45
	2.07
	116.0
	2.04
	114.0

	TSM01THg2
	387.009.F.01
	1.14
	4.81
	108.0
	4.63
	103.0


Figure A1-5. Control chart for MS/MSD sample analysis for Total Mercury. 


[image: image5.wmf]Matrix Spike Sample Analysis Total Mercury

control limit = 75-125% recovery

n = 10, mean recovery = 103.3%

50

70

90

110

130

150

TSM00THg1

TSM00THg2

TSM00THg3

TSM01THg1

TSM01THg2

Batch Identification

Percent Recovery


























_1158084177.xls
Chart4

		TSM00THg1		TSM00THg1		TSM00THg1		TSM00THg1		TSM00THg1

		TSM00THg2		TSM00THg2		TSM00THg2		TSM00THg2		TSM00THg2

		TSM00THg3		TSM00THg3		TSM00THg3		TSM00THg3		TSM00THg3

		TSM01THg1		TSM01THg1		TSM01THg1		TSM01THg1		TSM01THg1

		TSM01THg2		TSM01THg2		TSM01THg2		TSM01THg2		TSM01THg2



Result

Percent Recovery

Lower Control Limit

Upper Control Limit

True Value

Batch Identification

Percent Recovery

Certified Reference Material Total Mercury
NIST DORM-2 Dogfish Muscle certified value = 4.64 +/- 0.24 ppm
control limit = 75-125% recovery
n = 5, sample mean = 4.75 (102.4%)

4.71

101.5086206897

75

125

100

4.61

99.3534482759

75

125

100

4.46

96.1206896552

75

125

100

5.12

110.3448275862

75

125

100

4.85

104.525862069

75

125

100



CRM THg

		Batch Identification		Result		Percent Recovery		Lower Control Limit		Upper Control Limit		True Value

		TSM00THg1		4.71		101.5		75		125		100

		TSM00THg2		4.61		99.4		75		125		100

		TSM00THg3		4.46		96.1		75		125		100

		TSM01THg1		5.12		110.3		75		125		100

		TSM01THg2		4.85		104.5		75		125		100

		sample mean		4.75		102.4

		true value 4.64 +/-0.26





CRM THg

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0



Result

Percent Recovery

Lower Control Limit

Upper Control Limit

True Value

Batch Identification

Percent Recovery

Certified Reference Material Total Mercury
NIST DORM-2 Dogfish Muscle true value = 4.64 +/- 0.24 ppm
control limit = 75-125% recovery
n = 5, sample mean = 4.75 (102.4%)



Sheet2

		Batch Identification		Parent Sample Result		Replicate Sample Result		RPD





Sheet3

		






_1158084672.xls
Chart7

		TSM00THg1		TSM00THg1		75		125

		TSM00THg2		TSM00THg2		75		125

		TSM00THg3		TSM00THg3		75		125

		TSM01THg1		TSM01THg1		75		125

		TSM01THg2		TSM01THg2		75		125



Percent Recovery

Percent Recovery

Lower Control Limit

Upper Control Limit

Batch Identification

Percent Recovery

Matrix Spike Sample Analysis Total Mercury
control limit = 75-125% recovery
n = 10, mean recovery = 103.3%

101

103

93.4

100

96.5

98.3

116

114

108

103



CRM THg

		Batch Identification		Result		Percent Recovery		Lower Control Limit		Upper Control Limit		True Value

		TSM00THg1		4.71		101.5		75		125		100

		TSM00THg2		4.61		99.4		75		125		100

		TSM00THg3		4.46		96.1		75		125		100

		TSM01THg1		5.12		110.3		75		125		100

		TSM01THg2		4.85		104.5		75		125		100

		sample mean		4.75		102.4

		true value 4.64 +/-0.26





CRM THg

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0



Result

Percent Recovery

Lower Control Limit

Upper Control Limit

True Value

Batch Identification

Percent Recovery

Certified Reference Material Total Mercury
NIST DORM-2 Dogfish Muscle true value = 4.64 +/- 0.24 ppm
control limit = 75-125% recovery
n = 5, sample mean = 4.75 (102.4%)



REP THg

		Batch Identification		Sample Identification		Parent Sample Result		Replicate Sample Result		RPD		Control Limit

		TSM00THg1		049.002.F.00		5.26		5.53		5		25

		TSM00THg2		107.006.F.00		3.68		3.63		1.37		25

		TSM00THg3		384.006.F.00		0.914		0.861		5.97		25

		TSM01THg1		064.001.F.01		0.449		0.456		1.55		25

		TSM01THg2		387.009.F.01		1.14		1.11		2.67		25

		3.312





REP THg

		0		0

		0		0

		0		0

		0		0

		0		0



RPD

Control Limit

Batch Identification

Relative Percent Difference

Replicate Sample Analysis Total Mercury
control limit = 25% RPD
n = 10, mean RPD = 3.3



MSMSD THg

		Batch Identification		Sample Identification		Parent Sample Result		Matrix Spike Result		Percent Recovery		Matrix Spike Duplicate Result		Percent Recovery		Lower Control Limit		Upper Control Limit

		TSM00THg1		049.002.F.00		5.26		21.8		101.0		22.2		103.0		75		125

		TSM00THg2		107.006.F.00		3.68		14.0		93.4		14.7		100.0		75		125

		TSM00THg3		384.006.F.00		0.91		3.49		96.5		3.54		98.3		75		125

		TSM01THg1		064.001.F.01		0.45		2.07		116.0		2.04		114.0		75		125

		TSM01THg2		387.009.F.01		1.14		4.81		108.0		4.63		103.0		75		125

																		103.0

																		100.0

																		98.3

																		114.0

																		103.0

																		101.0

																		93.4

																		96.5

																		116.0

																		108.0

																		103.3





MSMSD THg

		0		0		0		0

		0		0		0		0

		0		0		0		0

		0		0		0		0

		0		0		0		0



Percent Recovery

Percent Recovery

Lower Control Limit

Upper Control Limit

Batch Identification

Percent Recovery

Matrix Spike Sample Analysis Total Mercury
control limit = 75-125% recovery
n = 10, mean recovery = 103.3%




_1158085532.xls
Chart5

		TSM00THg1		TSM00THg1

		TSM00THg2		TSM00THg2

		TSM00THg3		TSM00THg3

		TSM01THg1		TSM01THg1

		TSM01THg2		TSM01THg2



RPD

Control Limit

Batch Identification

Relative Percent Difference

Replicate Sample Analysis Total Mercury
control limit = 25% RPD, if sample result >10x MDL
n = 10, mean RPD = 3.3

5

25

1.37

25

5.97

25

1.55

25

2.67

25



CRM THg

		Batch Identification		Result		Percent Recovery		Lower Control Limit		Upper Control Limit		True Value

		TSM00THg1		4.71		101.5		75		125		100

		TSM00THg2		4.61		99.4		75		125		100

		TSM00THg3		4.46		96.1		75		125		100

		TSM01THg1		5.12		110.3		75		125		100

		TSM01THg2		4.85		104.5		75		125		100

		sample mean		4.75		102.4

		true value 4.64 +/-0.26





CRM THg

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0

		0		0		0		0		0



Result

Percent Recovery

Lower Control Limit

Upper Control Limit

True Value

Batch Identification

Percent Recovery

Certified Reference Material Total Mercury
NIST DORM-2 Dogfish Muscle true value = 4.64 +/- 0.24 ppm
control limit = 75-125% recovery
n = 5, sample mean = 4.75 (102.4%)



Sheet2

		Batch Identification		Sample Identification		Parent Sample Result		Replicate Sample Result		RPD		Control Limit

		TSM00THg1		049.002.F.00		5.26		5.53		5		25

		TSM00THg2		107.006.F.00		3.68		3.63		1.37		25

		TSM00THg3		384.006.F.00		0.914		0.861		5.97		25

		TSM01THg1		064.001.F.01		0.449		0.456		1.55		25

		TSM01THg2		387.009.F.01		1.14		1.11		2.67		25

		3.312





Sheet2

		0		0

		0		0

		0		0

		0		0

		0		0



RPD

Control Limit

Batch Identification

Relative Percent Difference

Replicate Sample Analysis Total Mercury
control limit = 25% RPD
n = 10, mean RPD = 3.3



Sheet3

		






_1158084473.xls
Chart2

		PCB 18		PCB 18		PCB 18		PCB 18		PCB 18		PCB 18		PCB 18		PCB 18		PCB 18

		PCB 28		PCB 28		PCB 28		PCB 28		PCB 28		PCB 28		PCB 28		PCB 28		PCB 28

		PCB 31		PCB 31		PCB 31		PCB 31		PCB 31		PCB 31		PCB 31		PCB 31		PCB 31

		PCB 99		PCB 99		PCB 99		PCB 99		PCB 99		PCB 99		PCB 99		PCB 99		PCB 99

		PCB 118		PCB 118		PCB 118		PCB 118		PCB 118		PCB 118		PCB 118		PCB 118		PCB 118

		PCB 128		PCB 128		PCB 128		PCB 128		PCB 128		PCB 128		PCB 128		PCB 128		PCB 128

		PCB 138		PCB 138		PCB 138		PCB 138		PCB 138		PCB 138		PCB 138		PCB 138		PCB 138

		PCB 149		PCB 149		PCB 149		PCB 149		PCB 149		PCB 149		PCB 149		PCB 149		PCB 149

		PCB 153		PCB 153		PCB 153		PCB 153		PCB 153		PCB 153		PCB 153		PCB 153		PCB 153

		PCB 180		PCB 180		PCB 180		PCB 180		PCB 180		PCB 180		PCB 180		PCB 180		PCB 180

		PCB 194		PCB 194		PCB 194		PCB 194		PCB 194		PCB 194		PCB 194		PCB 194		PCB 194

		PCB 195		PCB 195		PCB 195		PCB 195		PCB 195		PCB 195		PCB 195		PCB 195		PCB 195

		PCB 201		PCB 201		PCB 201		PCB 201		PCB 201		PCB 201		PCB 201		PCB 201		PCB 201

		PCB 203		PCB 203		PCB 203		PCB 203		PCB 203		PCB 203		PCB 203		PCB 203		PCB 203



PCB Congener

Percent Recovery

Matrix Spike Sample Analysis PCB Congeners
control limit = 50-150% recovery
n = 76

91.2894873925

87.9176661376

89.6278336977

95.7315409957

64.7448979592

67.4168797954

50

150

100

121.038201234

117.6162241296

104.9421332662

103.4075456727

95.306122449

95.3964194373

50

150

100

107.5865473681

88.8

91.2716051977

91.6592784769

103.0612244898

104.1432225064

50

150

100

110.3385592589

106.0476944923

97.123411688

93.9936300302

98.5204081633

93.2992327366

50

150

100

121

116

104.7448979592

96.8797953964

50

150

100

116.6603792212

126.2640734576

105.0409141523

104.0464870043

71.9897959184

72.0716112532

50

150

100

112

104

125.0510204082

100.716112532

50

150

100

115

113

97.3

94.2

121.9897959184

103.8874680307

50

150

100

126

112

136.9897959184

106.547314578

50

150

100

112

108

117.193877551

95.1918158568

50

150

100

108

112

90.4333995745

91.954434816

90.306122449

85.1662404092

50

150

100

110

109

121.2094867452

122.2903402713

93.2653061224

90.5882352941

50

150

100

110.7517317381

109.0608382314

100.3761367348

100.4322643857

97.193877551

90.0255754476

50

150

100

57

90.343772922

89.0753530177

78.992191585

106.1224489796

97.5959079284

50

150

100



CRM-PCB

		QA-Designated								95% CI

		Batch		PCB		Certified				Ranges				70-130% of the 95%

		Identification		Congener		Value		+/-		Lower		Upper		Confidence Interval				Result		Result		RPD

		TSM2000		28		7.91		0.9		7.01		8.81		4.91		11.45		7.45		7.78		4.3

		TSM2000		31		21.4		0.43		21.0		21.83		14.68		28.38		6.05		6.34		4.7

		TSM2000		44		11.8		0.64		11.16		12.44		7.81		16.17		11.3		10.9		3.9

		TSM2000		49		16.84		0.86		16.0		17.7		11.19		23.01		14.6		13.2		10.0

		TSM2000		52		17.7		2.8		14.9		20.5		10.43		26.65		17.8		16.3		8.7

		TSM2000		66		18.4		1.5		16.9		19.9		11.83		25.87		24.5		22.5		8.8

		TSM2000		87		10.2		0.29		9.91		10.49		6.94		13.64		12.3		11.9		3.2

		TSM2000		95		20.8		2.1		18.7		22.9		13.09		29.77		23.0		19.4		17.3

		TSM2000		99		18.84		0.44		18.4		19.28		12.88		25.06		20.2		19.0		6.0

		TSM2000		101		35.9		1.6		34.3		37.5		24.01		48.75		46.6		45.2		3.1

		TSM2000		105		10.8		0.45		10.35		11.25		7.25		14.63		7.28		12.8		54.7

		TSM2000		110		35.34		0.71		34.63		36.05		24.24		46.87		37.0		42.3		13.3

		TSM2000		118		35.1		1.0		34.1		36.1		23.87		46.93		40.0		41.7		4.1

		TSM2000		128		5.25		0.17		5.08		5.42		3.56		7.05		4.90		5.47		10.9

		TSM2000		138		35.7		1.5		34.2		37.2		23.94		48.36		46.5		46.5		0.1

		TSM2000		149		34.73		0.69		34.0		35.42		23.83		46.05		29.8		29.0		2.6

		TSM2000		151		10.92		0.25		10.67		11.17		7.47		14.52		11.1		11.3		1.8

		TSM2000		153		56.9		3.5		53.4		60.4		37.38		78.52		55.3		54.7		1.1

		TSM2000		156		1.97		0.11		1.86		2.08		1.30		2.70		1.99		2.09		4.7

		TSM2000		180		7.81		0.63		7.18		8.44		5.03		10.97		6.48		6.17		4.9

		TSM2000		183		5.25		0.15		5.10		5.40		3.57		7.02		4.65		4.14		11.6

		TSM2000		187		16.7		1.3		15.40		18.00		10.78		23.40		16.9		16.4		3.0

		TSM2001		28		7.91		0.9		7.01		8.81		4.91		11.45		5.30		NA		NA

		TSM2001		31		21.4		0.43		21.0		21.83		14.68		28.38		5.26		NA		NA

		TSM2001		44		11.8		0.64		11.16		12.44		7.81		16.17		11.0		NA		NA

		TSM2001		49		16.84		0.86		16.0		17.7		11.19		23.01		12.4		NA		NA

		TSM2001		52		17.7		2.8		14.9		20.5		10.43		26.65		15.5		NA		NA

		TSM2001		66		18.4		1.5		16.9		19.9		11.83		25.87		17.6		NA		NA

		TSM2001		87		10.2		0.29		9.91		10.49		6.94		13.64		10.4		NA		NA

		TSM2001		95		20.8		2.1		18.7		22.9		13.09		29.77		19.6		NA		NA

		TSM2001		99		18.84		0.44		18.4		19.28		12.88		25.06		18.3		NA		NA

		TSM2001		101		35.9		1.6		34.3		37.5		24.01		48.75		42.3		NA		NA

		TSM2001		105		10.8		0.45		10.35		11.25		7.25		14.63		9.19		NA		NA

		TSM2001		110		35.34		0.71		34.63		36.05		24.24		46.87		37.4		NA		NA

		TSM2001		118		35.1		1.0		34.1		36.1		23.87		46.93		36		NA		NA

		TSM2001		128		5.25		0.17		5.08		5.42		3.56		7.05		4.86		NA		NA

		TSM2001		138		35.7		1.5		34.2		37.2		23.94		48.36		46.6		NA		NA

		TSM2001		149		34.73		0.69		34.0		35.42		23.83		46.05		29.9		NA		NA

		TSM2001		151		10.92		0.25		10.67		11.17		7.47		14.52		9.54		NA		NA

		TSM2001		153		56.9		3.5		53.4		60.4		37.38		78.52		59.8		NA		NA

		TSM2001		156		1.97		0.11		1.86		2.08		1.30		2.70		1.76		NA		NA

		TSM2001		180		7.81		0.63		7.18		8.44		5.03		10.97		5.64		NA		NA

		TSM2001		183		5.25		0.15		5.10		5.40		3.57		7.02		4.43		NA		NA

		TSM2001		187		16.7		1.3		15.40		18.00		10.78		23.40		16.1		NA		NA





Blanks-PCB

		QA-Designated				Method Blank		Method Blank

		Batch Identification		PCB Congener		Sample Result		Sample Result

		TSM2000		18		ND		ND

		TSM2000		28		0.096		0.088

		TSM2000		31		0.099		0.091

		TSM2000		99		0.044		0.060

		TSM2000		118		0.212		0.237

		TSM2000		128		ND		ND

		TSM2000		138		0.112		0.129

		TSM2000		149		ND		0.044

		TSM2000		153		0.053		0.065

		TSM2000		180		ND		ND

		TSM2000		194		ND		ND

		TSM2000		195		ND		ND

		TSM2000		201		ND		ND

		TSM2000		203		ND		ND

		TSM2001		18		ND		NA

		TSM2001		28		<MDL		NA

		TSM2001		31		<MDL		NA

		TSM2001		99		<MDL		NA

		TSM2001		118		0.296		NA

		TSM2001		128		<MDL		NA

		TSM2001		138		0.198		NA

		TSM2001		149		<MDL		NA

		TSM2001		153		<MDL		NA

		TSM2001		180		ND		NA

		TSM2001		194		ND		NA

		TSM2001		195		ND		NA

		TSM2001		201		ND		NA

		TSM2001		203		ND		NA





REP-PCB

		QA-Designated																QA-Designated				TSM2000		TSM2000

		Batch		Sample				Parent		Replicate Sample				Sample				Batch								Control Limit

		Identification		Identification		PCB Congener		Sample Result		Result		RPD		<10xMDL				Identification		PCB Congener		RPD		RPD

		TSM2000		384.010.F.00		18		0.314		0.334		6.2		x				TSM2000		PCB 18						25

		TSM2000		384.010.F.00		28		1.17		1.18		0.2						TSM2000		PCB 28		0.2				25

		TSM2000		384.010.F.00		31		1.01		1.01		0.2						TSM2000		PCB 31		0.2				25

		TSM2000		384.010.F.00		99		6.26		6.16		1.6						TSM2000		PCB 99		1.6				25

		TSM2000		384.010.F.00		118		15.3		15.9		3.6						TSM2000		PCB 118		3.6		9.2		25

		TSM2000		384.010.F.00		128		3.89		4.95		24.0						TSM2000		PCB 128		24.0				25

		TSM2000		384.010.F.00		138		37.8		39.4		4.0						TSM2000		PCB 138		4.0		8.4		25

		TSM2000		384.010.F.00		149		11.2		11.3		1.4						TSM2000		PCB 149		1.4				25

		TSM2000		384.010.F.00		153		44.0		46.5		5.6						TSM2000		PCB 153		5.6		7.5		25

		TSM2000		384.010.F.00		180		21.6		22.5		4.1						TSM2000		PCB 180		4.1				25

		TSM2000		384.010.F.00		194		3.09		3.09		0.1						TSM2000		PCB 194		0.1				25

		TSM2000		384.010.F.00		195		0.596		0.497		18.1		x				TSM2000		PCB 195						25

		TSM2000		384.010.F.00		201		4.85		4.97		2.5						TSM2000		PCB 201		2.5				25

		TSM2000		384.010.F.00		203		3.51		3.57		1.6						TSM2000		PCB 203		1.6				25

		TSM2000		384.001.F.00		18		0.135		0.126		7.2		x

		TSM2000		384.001.F.00		28		0.288		0.302		5.0		x

		TSM2000		384.001.F.00		31		0.230		0.277		18.9		x

		TSM2000		384.001.F.00		99		0.479		0.509		6.0		x

		TSM2000		384.001.F.00		118		1.04		1.14		9.2

		TSM2000		384.001.F.00		128		0.240		0.264		9.7		x

		TSM2000		384.001.F.00		138		1.66		1.81		8.4

		TSM2000		384.001.F.00		149		0.752		0.818		8.4		x

		TSM2000		384.001.F.00		153		1.51		1.63		7.5

		TSM2000		384.001.F.00		180		0.662		0.696		4.9		x

		TSM2000		384.001.F.00		194		0.107		0.105		2.3		x

		TSM2000		384.001.F.00		195		ND		ND		ND

		TSM2000		384.001.F.00		201		0.174		0.173		0.8		x

		TSM2000		384.001.F.00		203		0.122		0.169		32.6		x

		TSM2001		013.002.F.01		18		ND		ND		ND		x

		TSM2001		013.002.F.01		28		0.102		0.276		92.1		x

		TSM2001		013.002.F.01		31		0.093		0.224		82.6		x

		TSM2001		013.002.F.01		99		0.235		0.273		15.0		x

		TSM2001		013.002.F.01		118		0.655		0.834		24.0		x

		TSM2001		013.002.F.01		128		0.113		0.112		0.9		x

		TSM2001		013.002.F.01		138		0.950		0.846		11.6		x

		TSM2001		013.002.F.01		149		0.298		0.287		3.8		x

		TSM2001		013.002.F.01		153		0.750		0.680		9.8		x

		TSM2001		013.002.F.01		180		0.251		0.238		5.3		x

		TSM2001		013.002.F.01		194		0.038		0.024		<MDL		x

		TSM2001		013.002.F.01		195		ND		ND		ND		x

		TSM2001		013.002.F.01		201		0.031		0.027		<MDL		x

		TSM2001		013.002.F.01		203		0.043		0.039		<MDL		x





REP-PCB

		0		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0



PCB Congener

Relative Percent Difference

Replicate Sample Analysis PCB Congeners
control limit = 25% if sample result 10x MDL
n = 15, mean RPD = 4.9



Sheet1

		QA-Designated		Parent

		Batch		Sample				MS		Percent		MSD		Percent								Percent		Percent

		Identification		Identification		PCB Congener		Result		Recovery		Result		Recovery						PCB Congener		Recovery		Recovery										Lower Control Limit		Upper Control Limit

		TSM2000		384.012.F.00		18		2.06		91.3		2.00		87.9						PCB 18		91.3		87.9		89.6		95.7		64.7		67.4		50		150		100.0

		TSM2000		384.012.F.00		28		3.23		121		3.18		118						PCB 28		121		118		105		103		95.3		95.4		50		150		100

		TSM2000		384.012.F.00		31		2.86		108		2.50		88.8						PCB 31		108		88.8		91.3		91.7		103		104		50		150		100

		TSM2000		384.012.F.00		99		6.23		110		6.16		106						PCB 99		110		106		97.1		94.0		98.5		93.3		50		150		100

		TSM2000		384.012.F.00		118		12.4		**		12.5		**						PCB 118						121		116		105		96.9		50		150		100

		TSM2000		384.012.F.00		128		4.96		117		5.16		126						PCB 128		117		126		105		104		72.0		72.1		50		150		100

		TSM2000		384.012.F.00		138		27.1		**		26.7		**						PCB 138						112		104		125		101		50		150		100

		TSM2000		384.012.F.00		149		8.29		115		8.26		113						PCB 149		115		113		97.3		94.2		122		104		50		150		100

		TSM2000		384.012.F.00		153		31.2		**		30.7		**						PCB 153						126		112		137		107		50		150		100

		TSM2000		384.012.F.00		180		17.1		**		17.0		**						PCB 180						112		108		117		95.2		50		150		100

		TSM2000		384.012.F.00		194		4.23		108		4.32		112						PCB 194		108		112		90.4		92.0		90.3		85.2		50		150		100

		TSM2000		384.012.F.00		195		2.98		110		2.97		109						PCB 195		110		109		121		122		93.3		90.6		50		150		100

		TSM2000		384.012.F.00		201		5.38		111		5.36		109						PCB 201		111		109		100		100		97.2		90.0		50		150		100

		TSM2000		384.012.F.00		203		3.46		57.0		4.14		90.3						PCB 203		57.0		90.3		89.1		79.0		106		97.6		50		150		100

		TSM2000		384.002.F.00		18		1.90		89.6		2.00		95.7

		TSM2000		384.002.F.00		28		2.59		105		2.54		103

		TSM2000		384.002.F.00		31		2.10		91.3		2.09		91.7

		TSM2000		384.002.F.00		99		3.65		97.1		3.57		94.0

		TSM2000		384.002.F.00		118		6.41		121		6.29		116

		TSM2000		384.002.F.00		128		2.97		105		2.93		104

		TSM2000		384.002.F.00		138		8.38		112		8.20		104

		TSM2000		384.002.F.00		149		3.32		97.3		3.24		94.2

		TSM2000		384.002.F.00		153		10.6		126		10.3		112

		TSM2000		384.002.F.00		180		5.35		112		5.25		108

		TSM2000		384.002.F.00		194		2.31		90.4		2.32		92.0

		TSM2000		384.002.F.00		195		2.59		121		2.59		122

		TSM2000		384.002.F.00		201		2.71		100		2.69		100

		TSM2000		384.002.F.00		203		2.40		89.1		2.18		79.0

		TSM2001		064.8.F.01		18		1.61		64.7		1.66		67.4

		TSM2001		064.8.F.01		28		1.95		95.3		1.95		95.4

		TSM2001		064.8.F.01		31		2.16		103		2.18		104

		TSM2001		064.8.F.01		99		2.60		98.5		2.49		93.3

		TSM2001		064.8.F.01		118		3.42		105		3.26		96.9

		TSM2001		064.8.F.01		128		1.92		72.0		1.92		72.1

		TSM2001		064.8.F.01		138		6.87		125		6.39		101

		TSM2001		064.8.F.01		149		4.65		122		4.29		104

		TSM2001		064.8.F.01		153		7.47		137		6.87		107

		TSM2001		064.8.F.01		180		5.28		117		4.84		95.2

		TSM2001		064.8.F.01		194		2.25		90.3		2.14		85.2

		TSM2001		064.8.F.01		195		2.03		93.3		1.98		90.6

		TSM2001		064.8.F.01		201		2.55		97.2		2.40		90.0

		TSM2001		064.8.F.01		203		2.65		106		2.48		97.6





Sheet1

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0

		0		0		0		0		0		0		0		0		0



PCB Congener

Percent Recovery

Matrix Spike Sample Analysis PCB Congeners
control limit = 50-150% recovery
n = 76




_1158084165.xls
Chart1

		PCB 18		PCB 18		PCB 18		PCB 18

		PCB 28		PCB 28		PCB 28		PCB 28

		PCB 31		PCB 31		PCB 31		PCB 31

		PCB 99		PCB 99		PCB 99		PCB 99

		PCB 118		PCB 118		PCB 118		PCB 118

		PCB 128		PCB 128		PCB 128		PCB 128

		PCB 138		PCB 138		PCB 138		PCB 138

		PCB 149		PCB 149		PCB 149		PCB 149

		PCB 153		PCB 153		PCB 153		PCB 153

		PCB 180		PCB 180		PCB 180		PCB 180

		PCB 194		PCB 194		PCB 194		PCB 194

		PCB 195		PCB 195		PCB 195		PCB 195

		PCB 201		PCB 201		PCB 201		PCB 201

		PCB 203		PCB 203		PCB 203		PCB 203



PCB Congener

Relative Percent Difference

Replicate Sample Analysis PCB Congeners
control limit = 25%, if sample result >10x MDL
n = 15, mean RPD = 4.9

25

0.15304964

25

0.2461571406

25

1.6290827676

25

3.5613389709

9.2351482192

25

23.9586835588

25

4.0406740784

8.3695608073

25

1.3579343642

25

5.5519930195

7.536600967

25

4.0978620119

25

0.1353867693

25

25

2.481430773

25

1.5510944296

25



CRM-PCB

		QA-Designated								95% CI

		Batch		PCB		Certified				Ranges				70-130% of the 95%

		Identification		Congener		Value		+/-		Lower		Upper		Confidence Interval				Result		Result		RPD

		TSM2000		28		7.91		0.9		7.01		8.81		4.91		11.45		7.45		7.78		4.3

		TSM2000		31		21.4		0.43		21.0		21.83		14.68		28.38		6.05		6.34		4.7

		TSM2000		44		11.8		0.64		11.16		12.44		7.81		16.17		11.3		10.9		3.9

		TSM2000		49		16.84		0.86		16.0		17.7		11.19		23.01		14.6		13.2		10.0

		TSM2000		52		17.7		2.8		14.9		20.5		10.43		26.65		17.8		16.3		8.7

		TSM2000		66		18.4		1.5		16.9		19.9		11.83		25.87		24.5		22.5		8.8

		TSM2000		87		10.2		0.29		9.91		10.49		6.94		13.64		12.3		11.9		3.2

		TSM2000		95		20.8		2.1		18.7		22.9		13.09		29.77		23.0		19.4		17.3

		TSM2000		99		18.84		0.44		18.4		19.28		12.88		25.06		20.2		19.0		6.0

		TSM2000		101		35.9		1.6		34.3		37.5		24.01		48.75		46.6		45.2		3.1

		TSM2000		105		10.8		0.45		10.35		11.25		7.25		14.63		7.28		12.8		54.7

		TSM2000		110		35.34		0.71		34.63		36.05		24.24		46.87		37.0		42.3		13.3

		TSM2000		118		35.1		1.0		34.1		36.1		23.87		46.93		40.0		41.7		4.1

		TSM2000		128		5.25		0.17		5.08		5.42		3.56		7.05		4.90		5.47		10.9

		TSM2000		138		35.7		1.5		34.2		37.2		23.94		48.36		46.5		46.5		0.1

		TSM2000		149		34.73		0.69		34.0		35.42		23.83		46.05		29.8		29.0		2.6

		TSM2000		151		10.92		0.25		10.67		11.17		7.47		14.52		11.1		11.3		1.8

		TSM2000		153		56.9		3.5		53.4		60.4		37.38		78.52		55.3		54.7		1.1

		TSM2000		156		1.97		0.11		1.86		2.08		1.30		2.70		1.99		2.09		4.7

		TSM2000		180		7.81		0.63		7.18		8.44		5.03		10.97		6.48		6.17		4.9

		TSM2000		183		5.25		0.15		5.10		5.40		3.57		7.02		4.65		4.14		11.6

		TSM2000		187		16.7		1.3		15.40		18.00		10.78		23.40		16.9		16.4		3.0

		TSM2001		28		7.91		0.9		7.01		8.81		4.91		11.45		5.30		NA		NA

		TSM2001		31		21.4		0.43		21.0		21.83		14.68		28.38		5.26		NA		NA

		TSM2001		44		11.8		0.64		11.16		12.44		7.81		16.17		11.0		NA		NA

		TSM2001		49		16.84		0.86		16.0		17.7		11.19		23.01		12.4		NA		NA

		TSM2001		52		17.7		2.8		14.9		20.5		10.43		26.65		15.5		NA		NA

		TSM2001		66		18.4		1.5		16.9		19.9		11.83		25.87		17.6		NA		NA

		TSM2001		87		10.2		0.29		9.91		10.49		6.94		13.64		10.4		NA		NA

		TSM2001		95		20.8		2.1		18.7		22.9		13.09		29.77		19.6		NA		NA

		TSM2001		99		18.84		0.44		18.4		19.28		12.88		25.06		18.3		NA		NA

		TSM2001		101		35.9		1.6		34.3		37.5		24.01		48.75		42.3		NA		NA

		TSM2001		105		10.8		0.45		10.35		11.25		7.25		14.63		9.19		NA		NA

		TSM2001		110		35.34		0.71		34.63		36.05		24.24		46.87		37.4		NA		NA

		TSM2001		118		35.1		1.0		34.1		36.1		23.87		46.93		36		NA		NA

		TSM2001		128		5.25		0.17		5.08		5.42		3.56		7.05		4.86		NA		NA

		TSM2001		138		35.7		1.5		34.2		37.2		23.94		48.36		46.6		NA		NA

		TSM2001		149		34.73		0.69		34.0		35.42		23.83		46.05		29.9		NA		NA

		TSM2001		151		10.92		0.25		10.67		11.17		7.47		14.52		9.54		NA		NA

		TSM2001		153		56.9		3.5		53.4		60.4		37.38		78.52		59.8		NA		NA

		TSM2001		156		1.97		0.11		1.86		2.08		1.30		2.70		1.76		NA		NA

		TSM2001		180		7.81		0.63		7.18		8.44		5.03		10.97		5.64		NA		NA

		TSM2001		183		5.25		0.15		5.10		5.40		3.57		7.02		4.43		NA		NA

		TSM2001		187		16.7		1.3		15.40		18.00		10.78		23.40		16.1		NA		NA





Blanks-PCB

		QA-Designated				Method Blank		Method Blank

		Batch Identification		PCB Congener		Sample Result		Sample Result

		TSM2000		18		ND		ND

		TSM2000		28		0.096		0.088

		TSM2000		31		0.099		0.091

		TSM2000		99		0.044		0.060

		TSM2000		118		0.212		0.237

		TSM2000		128		ND		ND

		TSM2000		138		0.112		0.129

		TSM2000		149		ND		0.044

		TSM2000		153		0.053		0.065

		TSM2000		180		ND		ND

		TSM2000		194		ND		ND

		TSM2000		195		ND		ND

		TSM2000		201		ND		ND

		TSM2000		203		ND		ND

		TSM2001		18		ND		NA

		TSM2001		28		<MDL		NA

		TSM2001		31		<MDL		NA

		TSM2001		99		<MDL		NA

		TSM2001		118		0.296		NA

		TSM2001		128		<MDL		NA

		TSM2001		138		0.198		NA

		TSM2001		149		<MDL		NA

		TSM2001		153		<MDL		NA

		TSM2001		180		ND		NA

		TSM2001		194		ND		NA

		TSM2001		195		ND		NA

		TSM2001		201		ND		NA

		TSM2001		203		ND		NA





REP-PCB

		QA-Designated														QA-Designated				TSM2000		TSM2000

		Batch				Parent		Replicate Sample				Sample				Batch								Control Limit

		Identification		PCB Congener		Sample Result		Result		RPD		<10xMDL				Identification		PCB Congener		RPD		RPD

		TSM2000		18		0.314		0.334		6.2		x				TSM2000		PCB 18						25

		TSM2000		28		1.17		1.18		0.2						TSM2000		PCB 28		0.2				25

		TSM2000		31		1.01		1.01		0.2						TSM2000		PCB 31		0.2				25

		TSM2000		99		6.26		6.16		1.6						TSM2000		PCB 99		1.6				25

		TSM2000		118		15.3		15.9		3.6						TSM2000		PCB 118		3.6		9.2		25

		TSM2000		128		3.89		4.95		24.0						TSM2000		PCB 128		24.0				25

		TSM2000		138		37.8		39.4		4.0						TSM2000		PCB 138		4.0		8.4		25

		TSM2000		149		11.2		11.3		1.4						TSM2000		PCB 149		1.4				25

		TSM2000		153		44.0		46.5		5.6						TSM2000		PCB 153		5.6		7.5		25

		TSM2000		180		21.6		22.5		4.1						TSM2000		PCB 180		4.1				25

		TSM2000		194		3.09		3.09		0.1						TSM2000		PCB 194		0.1				25

		TSM2000		195		0.596		0.497		18.1		x				TSM2000		PCB 195						25

		TSM2000		201		4.85		4.97		2.5						TSM2000		PCB 201		2.5				25

		TSM2000		203		3.51		3.57		1.6						TSM2000		PCB 203		1.6				25

		TSM2000		18		0.135		0.126		7.2		x

		TSM2000		28		0.288		0.302		5.0		x

		TSM2000		31		0.230		0.277		18.9		x

		TSM2000		99		0.479		0.509		6.0		x

		TSM2000		118		1.04		1.14		9.2

		TSM2000		128		0.240		0.264		9.7		x

		TSM2000		138		1.66		1.81		8.4

		TSM2000		149		0.752		0.818		8.4		x

		TSM2000		153		1.51		1.63		7.5

		TSM2000		180		0.662		0.696		4.9		x

		TSM2000		194		0.107		0.105		2.3		x

		TSM2000		195		ND		ND		ND

		TSM2000		201		0.174		0.173		0.8		x

		TSM2000		203		0.122		0.169		32.6		x

		TSM2001		18		ND		ND		ND		x

		TSM2001		28		0.102		0.276		92.1		x

		TSM2001		31		0.093		0.224		82.6		x

		TSM2001		99		0.235		0.273		15.0		x

		TSM2001		118		0.655		0.834		24.0		x

		TSM2001		128		0.113		0.112		0.9		x

		TSM2001		138		0.950		0.846		11.6		x

		TSM2001		149		0.298		0.287		3.8		x

		TSM2001		153		0.750		0.680		9.8		x

		TSM2001		180		0.251		0.238		5.3		x

		TSM2001		194		0.038		0.024		<MDL		x

		TSM2001		195		ND		ND		ND		x

		TSM2001		201		0.031		0.027		<MDL		x

		TSM2001		203		0.043		0.039		<MDL		x





REP-PCB

		0		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0

		0		0		0



PCB Congener

Relative Percent Difference

Replicate Sample Analysis PCB Congeners
control limit = 25% if sample result 10x MDL
n = 15, mean RPD = 4.9




