## Conventional constituent in H2O
Data Acceptability Criteria for Conventional Constituents in Water
Sample Type | Objective | Frequency of Analysis | Recommended Control Limits | Recommended Corrective Action
External Calibration
Calibration Standards (3-5 standards over the expected range of sample target analyte conc., with the lowest conc. Std at or near the MDL). | Full calibration: Establish relationship between instrument response and target analyte conc. | Follow manufacturer's or procedures in specific analytical protocols. A min. 3 point calib.: Each set up, major disruption, and when routine calib check exceeds specific control limits. | Linear regression, r>0.995. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Calibration Verification
Calibration Check Standards (minimum of one mid-range standard prepared independently from initial calibration standards). | Verify calibration. | After initial calibration or recalibration. Every 20 samples. | %Recovery = 80 -120% | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Method Detection Limit Determination (MDL)
Spiked matrix samples (analyte-free water samples to which known amounts of target analytes have been added; one spike for each target analyte at 3-10 times the estimated MDL). | Establish or confirm MDL for analyte of interest. | Seven replicate analyses prior to use of method. Re-evaluation of MDL annually. | Determined by program manager. | Redetermine MDL.
Accuracy and Precision Assessment
Reference materials (SRMs or CRMs, covering the range of expected target analyte conc). | Assess method performance (initial method validation and routine accuracy assessment). | Method validation: As many as required to assess accuracy and precision of method before routine analysis of samples. Routine accuracy assessment: one (preferably blind) per 20 samples or one batch. | Measured value <95% confidence intervals, if certified. Otherwise, %Recovery = 80-120% | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Matrix spikes (field water samples to which known amounts of target analytes have been added: 5 times the concentration of analyte of interest or 10 times MDL). | Assess matrix effects and accuracy (%R) routinely. | One per 20 samples or one per batch, whichever is more frequent. | %Recovery = 80-120% or Control Limits based on 3x the standard deviation of laboratory's actual method recoveries. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data. Zero percent recovery requires rejection of all suspect data.
Matrix spike replicates (replicate aliquots of matrix spike samples; 5 times the concentration of analyte of interest or 10 times MDL). | Assess method precision routinely. | One duplicate per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | Determine cause, take appropriate corrective action. Recalibrate & reanalyze all suspect samples or flag all suspect data.
Laboratory Duplicate | Assess method precision | One per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | Determine cause, take appropriate corrective action. Recalibrate & reanalyze all suspect samples or flag all suspect data.
Field Replicate (replicate aliquots of water field samples). | Assess method precision routinely. Assess total variability (i.e., population variability, field or sampling variability, & analytical method variability.) | 5% annual rate (5% of total number of field samples per analytical procedure per year, rounded up to nearest whole number). | RPD <25% for duplicates. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Contamination Assessment
Laboratory Blanks (method, processing, bottle, reagent). | Assess contamination from equipment, reagents, etc. | One method blank per 20 samples or one per batch, whichever is more frequent. At least one bottle blank per batch. One reagent blank prior to use of a new batch of reagent and whenever method blank exceeds control limits. | Blanks <MDL for target analyte. | Determine cause of problem (e.g., contaminated reagents, equipment), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
Field Blanks, Travel Blanks, Equipment Blanks. | Assess contamination from equipment, from air, from surrounding environment, etc. | Water DOC must have field blanks and travel blanks analyzed at 5% rate. All else: random performance evaluation during field audit; field blanks <MDL for analyte of interest. If acceptable performance, no field blanks required until next field audit. If non-acceptable, 5% field blanks must be conducted until next field audit. | Blanks <MDL for target analyte. | Determine cause of problem (e.g., equipment contamination, improper cleaning, exposure to airborne contaminants, etc.), remove sources of contamination, & reanalyze all suspect samples or flag all suspect data.
External QA Assessment
Accuracy-based performance evaluation samples submitted to new laboratories by SWAMP QA Program. | Initial demonstration of laboratory capability. | Once prior to routine analysis of field samples. | Determined by study manager. | Determine cause of problem and reanalyze sample. Do not begin analysis of field samples until laboratory initial capability is clearly demonstrated.
Mandatory interlaboratory exercises overseen by 3rd party external ("referee") SWAMP QA Program officials for all SWAMP participant laboratories. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
Voluntary, but encouraged, participation in NOAA-NIST intercalibration studies & CA-ELAP annual performance evaluations, as appropriate. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
General Provisions
Acceptable Data Set: CCV Recoveries must be within control limits, & either SRM or Spiked Matrix recoveries must also be within control limits.

## VOA's in Water
Data Acceptability Criteria for Volatile Organic Analytes (including VOC's, MTBE, BTEX) in Water
Sample Type | Objective | Frequency of Analysis | Recommended Control Limits | Recommended Corrective Action
Internal Calibration Method
Calibration Standard (minimum of 5 levels over the expected range of sample target analyte conc., with the lowest conc. Std at or near the MDL). | Full calibration: Establish relationship between instrument response and target analyte conc. | Follow manufacturer's or procedures in specific analytical protocols. A min. 5 point calib.: Each set up, major disruption, and when routine calib check exceeds specific control limits. | % RSD for CCCs <30%. RF for SPCCs >0.1 except 1,1,2,2-tetrachloroethane which is 0.3. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Calibration Verification
Calibration Check Standard (minimum of one mid-range standard prepared independently from initial calibration standards). | Verify calibration. | Mid-level calibration run every 12 hours. | RF for SPCCs same as inititial calibration. RF of CCCs must be <20% difference from initial calibration. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Method Detection Limit Determination
Spiked matrix sample (analyte-free water samples to which known amounts of target analytes have been added; one spike for each target analyte at 3-10 times the estimated MDL). | Establish or confirm MDL for analyte of interest. | Seven replicate analyses prior to use of method. Re-evaluation of MDL annually. | Determined by program manager. | Redetermine MDL.
Accuracy and Precision Assessment
Reference material (SRMs or CRMs, covering the range of expected target analyte conc). | Assess method performance (initial method validation and routine accuracy assessment). | Method validation: As many as required to assess accuracy and precision of method before routine analysis of samples. Routine accuracy assessment: one (preferably blind) per 20 samples or one batch. | Measured value <95% confidence intervals, if certified. Otherwise, %Recovery = 50-150%. | Failure of any of the accuracy and precision control limits require the following: Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Matrix spike (field water samples to which known amounts of target analytes have been added: 5 times the concentration of the analyte of interest or 10 times the MDL). | Assess matrix effects and accuracy (%Recovery) routinely. | One per 20 samples or one per batch, whichever is more frequent. | %Recovery = 50-150% or Control Limits based on 3x the standard deviation of laboratory's actual method recoveries. | See Reference Materials Corrective Action. Zero percent recovery requires rejection of all suspect data.
Matrix spike replicate (replicate aliquots of matrix spike samples; 5 times the concentration of the analyte of interest or 10 times the MDL). | Assess method precision routinely. | One duplicate per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | See Reference Materials Corrective Action.
Field Replicate (replicate aliquots of water field samples). | Assess method precision routinely. Assess total variability (i.e., population variability, field or sampling variability, and analytical method variability). | 5% annual rate (5% of total number of field samples per analytical procedure per year, rounded up to nearest whole number). | RPD <25% for duplicates. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Contamination Assessment
Laboratory Blank (method, processing, bottle, reagent). | Assess contamination form equipment, reagents, etc. | One method blank per 20 samples or one per batch, whichever is more frequent. | Blanks <MDL for target analyte. | Determine cause of problem (e.g., contaminated reagents, equipment), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
Field Blanks, Travel Blanks, Equipment Blanks. | Assess contamination from equipment, from air, from surrounding environment, etc. | One field blank and one travel blank is required per 20 (or less) field samples collected for volatile organic analytes in water for each sampling "event". | Blanks <MDL for target analyte. | Determine cause of problem (e.g., equipment contamination, improper cleaning, exposure to airborne contaminants, etc.), remove sources of contamination, & reanalyze all suspect samples or flag all suspect data.
Routine Monitoring of Method Performance for Organic Analysis
Surrogate Spike (Prepared from chemicals of similar structure to target analytes or isotopically labelled target analyte). | Assess method performance and estimate recovery of target analytes analyzed by GC or GC/MS. Determine RRF's of organic target analytes quanitated by isotope dilution techniques. | Performed in every calibration standard, field sample, and blank analyzed for organics by GC or istope dilution GC-MS; added to samples prior to extraction. | Determined by program manager. | Determine cause of problem (e.g., incomplete extraction or digestion, contamination, inaccurate preparation of internal standard), take appropriate corrective action, and reanalyze all suspect samples or flag all suspect data.
External QA Assessment
Accuracy-based performance evaluation samples submitted to new laboratories by SWAMP QA Program. | Initial demonstration of laboratory capability. | Once prior to routine analysis of field samples. | Determined by study manager. | Determine cause of problem and reanalyze sample. Do not begin analysis of field samples until laboratory initial capability is clearly demonstrated.
Mandatory interlaboratory exercises overseen by 3rd party external ("referee") SWAMP QA Program officials for all SWAMP participant laboratories. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
Voluntary, but encouraged, participation in NOAA-NIST intercalibration studies and CA-ELAP annual performance evaluations, as appropriate. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
General Provisions
Acceptable Data Set: CCV Recoveries must be within control limits, & either SRM or Spiked Matrix recoveries must also be within control limits.

## Organics (non-VOC's) inWater
Data Acceptability Criteria for Synthetic Organic Compounds in Water
(non-volatiles) in water: PCB's, PAH's, Pesticides
Sample Type | Objective | Frequency of Analysis | Recommended Control Limits | Recommended Corrective Action
Internal/External Calibration
Calibration Standards (3-5 standards over the expected range of sample target analyte conc., with the lowest conc. Std at or near the MDL). | Full calibration: Establish relationship between instrument response and target analyte conc. | Follow manufacturer's or procedures in specific analytical protocols. A min. 3 point calib.: Each set up, major disruption, and when routine calib check exceeds specific control limits. | Linear regression, r>0.995 or %RSD <10%. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Calibration Verification
Calibration Check Standards (minimum of one mid-range standard prepared independently from initial calibration standards). | Verify calibration. | After initial calibration or recalibration. Every 10 samples. | %R = 85-115%. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Method Detection Limit Determination
Spiked matrix samples (analyte-free water samples to which known amounts of target analytes have been added; one spike for each target analyte at 3-10 times the estimated MDL). | Establish or confirm MDL for analyte of interest. | Seven replicate analyses prior to use of method, or 5 replicates at 3 concentrations with 1 level close to estimated RL. Re-evaluation of MDL annually. | Determined by program manager. | Redetermine MDL.
Accuracy and Precision Assessment
Reference materials (SRMs or CRMs, covering the range of expected target analyte conc) if available. | Assess method performance (initial method validation and routine accuracy assessment). | Method validation: As many as required to assess accuracy and precision of method before routine analysis of samples. Routine accuracy assessment: one (preferably blind) per 20 samples or one batch. | Measured value <95% confidence intervals, if certified. Otherwise, %Recovery = 50-150%. | Failure of any of the accuracy and precision control limits require the following: Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Matrix spikes (field water samples to which known amounts of target analytes have been added: 5 times the concentration of the analyte of interest or 10 times the MDL). | Assess matrix effects and accuracy (%Recovery) routinely. | One per 20 samples or one per batch, whichever is more frequent. | %Recovery = 50-150% or Control Limits based on 3x the standard deviation of laboratory's actual method recoveries. | See Reference Materials Corrective Action. Zero percent recovery requires rejection of all suspect data.
Matrix spike replicates (replicate aliquots of matrix spike samples; 5 times the concentration of the analyte of interest or 10 times the MDL). | Assess method precision routinely. | One duplicate per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | See Reference Materials Corrective Action.
Field Replicate (replicate aliquots of water field samples). | Assess method precision routinely. Assess total variability (i.e., population variability, field or sampling variability, and analytical method variability.) | 5% annual rate (5% of total number of field samples per analytical procedure per year, rounded up to nearest whole number). | RPD <25% for duplicates. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Contamination Assessment
Laboratory Blanks (method, processing, bottle, reagent). | Assess contamination from equipment, reagents, etc. | One method blank per 20 samples or one per batch, whichever is more frequent. At least one bottle blank per batch. One reagent blank prior to use of a new batch of reagent and whenever method blank exceeds control limits. | Blanks <MDL for target analytes. | Determine cause of problem (e.g., contaminated reagents, equipment), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
Field Blanks, Travel Blanks, Equipment Blanks. | Assess contamination from equipment, from air, from surrounding environment, etc. | Random performance evaluation during field audit; field blanks <MDL for analyte of interest. If acceptable performance, no field blanks required until next field audit. If non-acceptable, 5% field blanks must be conducted during the year until next field audit. | Blanks <MDL for target analytes. | Determine cause of problem (e.g., equipment contamination, improper cleaning, exposure to airborne contaminants, etc.), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
Routine Monitoring of Method Performance for Organic Analysis
Surrogate Spikes (Prepared from chemicals of similar structure to target analytes not in target list or isotopically labelled target analyte). | Assess method performance and estimate recovery of target analytes analyzed by GC or GC/MS. Determine RRF's of organic target analytes quanitated by isotope dilution techniques. | In every calibration standard, sample, and blank analyzed for organics by GC or istope dilution GC-MS; added to samples prior to extraction. | Determined by program manager. | Determine cause of problem (e.g., incomplete extraction or digestion, contamination, inaccurate preparation of internal standard), take appropriate corrective action, and reanalyze all suspect samples or flag all suspect data.
External QA Assessment
Accuracy-based performance evaluation samples submitted to new laboratories by SWAMP QA Program. | Initial demonstration of laboratory capability. | Once prior to routine analysis of field samples. | Determined by study manager. | Determine cause of problem and reanalyze sample. Do not begin analysis of field samples until laboratory initial capability is clearly demonstrated.
Mandatory interlaboratory exercises overseen by 3rd party external ("referee") SWAMP QA Program officials for all SWAMP participant laboratories. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
Voluntary, but encouraged, participation in NOAA-NIST intercalibration studies and CA-ELAP annual performance evaluations, as appropriate. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
General Provisions
Acceptable Data Set: CCV Recoveries must be within control limits, & either SRM or Spiked Matrix recoveries must also be within control limits.

## Trace metals incl Hg in Water
Data Acceptability Criteria for Trace Metals in Water, including Mercury
Sample Type | Objective | Frequency of Analysis | Recommended Control Limits | Recommended Corrective Action
External Calibration
Calibration Standards (3-5 standards over the expected range of sample target analyte conc., with the lowest conc. Std at or near the MDL). | Full calibration: Establish relationship between instrument response and target analyte conc. | Follow manufacturer's or procedures in specific analytical protocols. A min. 3 point calib.: Each set up, major disruption, and when routine calib check exceeds specific control limits. | Linear regression, r>0.995. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Calibration Verification
Calibration Check Standards (minimum of one mid-range standard prepared independently from initial calibration standards: an instrument internal standard must be added to each calib. check std. when internal std. calib. is being used). | Verify calibration. | After initial calibration or recalibration. Every 10 samples, | Mercury %R = 80-120%, all other metals %R = 90-110%. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Method Detection Limit Determination
Spiked matrix samples (analyte-free water samples to which known amounts of target analytes have been added; one spike for each target analyte at 3-10 times the estimated MDL). | Establish or confirm MDL for analyte of interest. | Seven replicate analyses prior to use of method. Reevaluation of MDL annually. | Determined by program manager. | Redetermine MDL.
Accuracy and Precision Assessment
Reference materials (SRMs or CRMs, covering the range of expected target analyte conc.). | Assess method performance (initial method validation and routine accuracy assessment). | Method validation: As many as required to assess accuracy and precision of method before routine analysis of samples. Routine accuracy assessment: one (preferably blind) per 20 samples or one batch. | Method validation and Routine accuracy assessment: %R = 75-125%. | If matrix spikes are in control then proceed. If not, determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Matrix spikes (field water samples to which known amounts of target analytes have been added: at least 5 times the concentration of the analyte of interest or 10 times the MDL). | Assess matrix effects and accuracy (%Recovery) routinely. | One per 20 samples or one per batch, whichever is more frequent. | %R = 75-125% | If SRMs are in control then proceed. If not, determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data. Zero percent recovery requires rejection of all suspect data.
Matrix spike replicates (replicate aliquots of matrix spike samples; at least 5 times the concentration of the analyte of interest or 10 times the MDL). | Assess method precision routinely. | One duplicate per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Laboratory Duplicate | Assess method precision | One per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | Determine cause, take appropriate corrective action. Recalibrate & reanalyze all suspect samples or flag all suspect data.
Field Replicate (replicate aliquots of water field samples). | Assess method precision routinely. Assess total variability (i.e., population variability, field or sampling variability, and analytical method variability.) | 5% annual rate (5% of total number of field samples per analytical procedure per year, rounded up to nearest whole number). | RPD <25% for duplicates. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Contamination Assessment
Laboratory Blanks (method, processing, bottle, filters, reagent). | Assess contamination from equipment, reagents, etc. | One method blank per 20 samples or one per batch, whichever is more frequent. At least one bottle blank per batch. One reagent blank prior to use of a new batch of reagent and whenever method blank exceeds control limits. One filter blank per new lot. | Blanks<MDL for target analytes. | Determine cause of problem (e.g., contaminated reagents, equipment), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
Field Blanks, Travel Blanks, Equipment Blanks. | Assess contamination from equipment, from air, from surrounding environment, etc. | One field equipment blank is required per 20 (or less) field water samples collected for trace metals analysis, including mercury. Pre-cleaned C-flex tubing used for each site w/peristaltic pump for depth-integration; or pre-cleaned syringe and pencil filter used at each site otherwise. | Blanks<MDL for target analytes. | Determine cause of problem (e.g., contaminated preservatives, equipment contamination, improper cleaning, exposure to airborne contaminants, etc.), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
External QA Assessment
Accuracy-based performance evaluation samples submitted to new laboratories by SWAMP QA Program. | Initial demonstration of laboratory capability. | Once prior to routine analysis of field samples. | Determined by study manager. | Determine cause of problem and reanalyze sample. Do not begin analysis of field samples until laboratory initial capability is clearly demonstrated.
Mandatory interlaboratory exercises overseen by 3rd party external ("referee") SWAMP QA Program officials for all SWAMP participant laboratories. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
Voluntary, but encouraged, participation in NOAA-NIST intercalibration studies and CA-ELAP annual performance evaluations, as appropriate. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
General Provisions
Acceptable Data Set: CCV Recoveries must be within control limits, & either SRM or Spiked Matrix recoveries must also be within control limits.

## Mono-Methyl Mercury in Water
Data Acceptability Criteria for Monomethyl Mercury in Water
Sample Type | Objective | Frequency of Analysis | Recommended Control Limits | Recommended Corrective Action
External Calibration
Calibration Standards (3-5 standards over the expected range of sample target analyte conc., with the lowest conc. Std at or near the MDL). | Full calibration: Establish relationship between instrument response and target analyte conc. | Follow manufacturer's or procedures in specific analytical protocols. A min., 3 point calib. Each set up, major disruption, and when routine calib check exceeds specific control limits. | Linear regression, r>0.995. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Calibration Verification
Calibration Check Standards (minimum of one mid-range standard prepared independently from initial calibration standards: an instrument internal standard must be added to each calib. check std. when internal std. calib. is being used). | Verify calibration. | After initial calibration or recalibration. Every 10 samples. | %R = 80-120%. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Method Detection Limit Determination
Spiked matrix samples (analyte-free water samples to which known amounts of target analytes have been added; one spike for each target analyte at 3-10 times the estimated MDL). | Establish or confirm MDL for analyte of interest. | Seven replicate analyses prior to use of method. Re-evaluation of MDL annually. | Determined by program manager. | Redetermine MDL.
Accuracy and Precision Assessment
Reference materials (SRMs or CRMs, covering the range of expected target analyte conc.). | Assess method performance (initial method validation and routine accuracy assessment). | Method validation: As many as required to assess accuracy & precision of method before routine analysis of samples. Routine accuracy assessment: one (preferably blind) per 20 samples or one batch. | Method validation and Routine accuracy assessment: %R = 70-130%. | If matrix spikes are in control then proceed. If not, determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Matrix spikes (field water samples to which known amounts of target analytes have been added: 5 times the concentration of the analyte of interest or 10 times the MDL). | Assess matrix effects and accuracy (%Recovery) routinely. | One per 20 samples or one per batch, whichever is more frequent. | %R = 70-130%. | If SRMs are in control then proceed. If not, determine cause & take appropriate corrective action. Recalibrate & reanalyze all suspect samples or flag all suspect data. Zero percent recovery requires rejection of all suspect data.
Matrix spike replicates (replicate aliquots of matrix spike samples; 5 times the concentration of analyte of interest or 10 times MDL). | Assess method precision routinely. | One duplicate per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Laboratory Duplicate | Assess method precision | One per 20 samples or one per batch, whichever is more frequent. | RPD <25% for duplicates. | Determine cause, take appropriate corrective action. Recalibrate & reanalyze all suspect samples or flag all suspect data.
Field Replicate (replicate aliquots of water field samples). | Assess method precision routinely. Assess total variability (i.e., population variability, field or sampling variability, and analytical method variability). | 5% annual rate (5% of total number of field samples per analytical procedure per year, rounded up to nearest whole number). | RPD <25% for duplicates. | Determine cause and take appropriate corrective action. Recalibrate and reanalyze all suspect samples or flag all suspect data.
Contamination Assessment
Laboratory Blanks (method, processing, bottle, reagent). | Assess contamination from equipment, reagents, etc. | Three method blank per 20 samples or one per batch, whichever is more frequent. At least one bottle blank per batch. One reagent blank prior to use of a new batch of reagent and whenever method blank exceeds control limits. | Blanks<MDL for target analytes. | Determine cause of problem (e.g., contaminated reagents, equipment), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
Field Blanks, Travel Blanks, Equipment Blanks. | Assess contamination from equipment, from air, from surrounding environment, etc. | One field equipment blank is required per 20 (or less) field water samples collected for trace metals analysis, including mercury. Pre-cleaned C-flex tubing used for each site w/peristaltic pump for depth-integration; or pre-cleaned syringe and pencil filter used at each site otherwise. | Blanks<MDL for target analytes. | Determine cause of problem (e.g., equipment contamination, improper cleaning, exposure to airborne contaminants, etc.), remove sources of contamination, and reanalyze all suspect samples or flag all suspect data.
External QA Assessment
Accuracy-based performance evaluation samples submitted to new laboratories by SWAMP QA Program. | Initial demonstration of laboratory capability. | Once prior to routine analysis of field samples. | Determined by study manager. | Determine cause of problem and reanalyze sample. Do not begin analysis of field samples until laboratory initial capability is clearly demonstrated.
Mandatory interlaboratory exercises overseen by 3rd party external ("referee") SWAMP QA Program officials for all SWAMP participant laboratories. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
Voluntary, but encouraged, participation in NOAA-NIST intercalibration studies and CA-ELAP annual performance evaluations, as appropriate. | Ongoing demonstration of laboratory capability. | One exercise per year. | Determined by study manager. | Determine cause of problem and reanalyze sample. Further corrective action to be determined by QA manager.
General Provisions
Acceptable Data Set: CCV Recoveries must be within control limits, & either SRM or Spiked Matrix recoveries must also be within control limits.