| YSI MULTIPROBE CALIBRATION/MAINTENANCE LOG |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Calibration |  |  |  |  |  |  |
| Date: | Time: |  | Employee Name: |  |  |  |
| Battery Voltage: |  |  | Sonde Type and Serial No. |  |  |  |
| Function |  | Temp. of Standard | Value of Standard | Initial Reading | Calibrated To | Comments |
| Specific Conductance (high) ~100 mS/cm |  |  |  |  |  |  |
| Specific Conductance (med) ~10 mS/cm |  |  |  |  |  |  |
| Specific Conductance (low) ~1 mS/cm |  |  |  |  |  |  |
| Conductivity Cell Constant (see *) |  |  |  |  |  | Range 4.6 to 5.45 |
| pH calibrated ( at pH 7) |  |  |  |  |  |  |
| pH mV for pH 7 solution |  |  |  |  |  | Range 0 to 40 mV |
| pH calibrated (at pH 4 10 – circle one) |  |  |  |  |  |  |
| pH mV for pH 10 pH mV for pH 4 |  |  |  |  |  | Range: -160 to 40 mV Range: +160 to 40 mV |
| Dissolved Oxygen (% sat) |  |  |  |  |  |  |
| Dissolved Oxygen Charge |  |  |  |  |  | Range 25 to 75 |
| Dissolved Oxygen Gain (see *) |  |  |  |  |  | Range 0.8 to 1.7 |
| DATA NEEDED FOR DISSOLVED OXYGEN CALIBRATION |  |  |  |  |  |  |
| Altitude (A )=______________feet above msl |  |  | Barometric Pressure _________ inches |  |  |  |
| Barometric Pressure (BP) Options |  |  | Barometirc Pressure Formulas |  |  |  |
| Barometer |  |  | Barometric Pressure (inches) ________ x 25.4 = BP ________mm |  |  |  |
| From Local Source After Correction (CBP) |  |  | BP _________ mm = CBP _______mm - 2.5 (altitude ____/100) |  |  |  |
| Estimated From Altitude Only |  |  | BP _________ mm= 760 mm - 2.5 (altitude _____/100) |  |  |  |
| Deployment Checklist (required for data logging only) |  |  |  |  |  |  |
| Logging Interval: Yes No |  |  | SDI-12 Autosleep enabled: Yes No |  |  |  |
| DO warm-up time: |  |  | Battery Volts in Sonde (days): |  |  |  |
| RS 232 Autosleep enabled: Yes No |  |  | Available Memory in Sonde (days): |  |  |  |
| Post-Calibration |  |  |  |  |  |  |
| Date: | Time: |  | Employee Name: |  |  |  |
| Battery Voltage: |  |  | Sonde Type and Serial No. |  |  |  |
| Function |  | Temp. of Standard | Value of Standard | Initial Reading | Comments |  |
| Specific Conductance |  |  |  |  |  |  |
| pH calibrated (~7) |  |  |  |  |  |  |
| pH slope (~ 4/10) |  |  |  |  |  |  |
| Dissolved Oxygen |  |  |  |  |  |  |
| Check previous maintenance and use; do the following before calibration: |  |  |  |  |  |  |
| Brush conductivity cell. Must be cleaned within the last two months or once every 15 field trips |  |  |  | Date: | Name/Comments: |  |
| Inspect D.O. membrane for nicks or bubbles. Must be changed within last six months or once every 15 field trips. |  |  |  | Date: | Name/Comments: |  |
| Comments: |  |  |  |  |  |  |