Surface Water Ambient Monitoring Program (SWAMP)

# Suspended Sediment and Turbidity Monitoring

# Freshwater Creek and Elk River

STANDARD OPERATING PROCEDURE 05

TURBIDITY THRESHOLD SAMPLING

Prepared for:

Salmon Forever

Arcata, CA

Prepared by:

Rand Eads

Surface Water Specialist

RiverMetrics LLC

PO Box 507

Lafayette, OR 97127

Version 1.0

February 2008

1. Procedures

a. Scope

This document, and the attached Turbidity Threshold Sampling Manual, provides detailed information to implement suspended sediment sampling at the gage locations in Freshwater Creek and Elk River.  The manual provides systematic instructions for operating the data logger program, servicing the equipment, and solving field problems.  Other Standard Operating Procedures (SOP 02, SOP 03, and SOP 04) provide additional technical information pertaining to depth-integrating sampling, discharge measurements, and field instrumentation.

b. Method Summary

The Turbidity Threshold Sampling Manual was originally developed by staff at Redwood Sciences Laboratory, USDA Forest Service.  The attached version has been updated, by RiverMetrics LLC, to be compatible with the newest version LoggerNet interface software.  The Turbidity Threshold Sampling method was developed to provide an effective and accurate method for developing sediment loads.  This method has gained wide acceptance and has been in use for more than 10 years.

Traditional methods for determining the frequency of suspended sediment sample collection often rely on measurements, such as water discharge, that are not well correlated to sediment concentration.  Stream power is generally not a good predictor of sediment concentration for rivers that transport the bulk of their load as fines, due to the highly variable routing of sediment to the channel from hillslopes, roads, and landslides.  A method, such as Turbidity Threshold Sampling, that employs a parameter well correlated to suspended sediment concentration, can improve sampling efficiency by collecting samples that are distributed over a range of rising and falling concentrations.  The resulting set of samples can be used to estimate sediment loads by establishing a relationship between concentration and turbidity for any sampled period and applying it to the continuous turbidity data.  All river systems, particularly smaller watersheds that respond very quickly to rainfall, benefit from automated data collection.

A data logger, under direction of the Turbidity Threshold Sampling program, collects stage and turbidity data at 10 or 15-minute intervals, depending on the drainage size, then triggers an automatic pumping sampler to collect a sample whenever the turbidity passes selected thresholds, if the stage is above a specified minimum and temperature is above freezing.  The program distinguishes rising from falling conditions and uses different sets of turbidity thresholds for each.  When the measured turbidity exceeds the calibrated range of the sensor, the program collects pumped samples at a fixed rate specified by the operator.  At each recording interval, the program records the median of 60-100 high-frequency turbidity readings.  To avoid over-sampling, the program logic attempts to distinguish brief spikes caused by fouling from true rises in turbidity.  Future improvements in program logic will provide a mechanism to limit the number of occurrences and frequency at which thresholds can be triggered without user intervention.

The turbidity probe, mounted inside of a protective housing, and the sampler intake, are usually attached to the end of an articulated sampling boom.  Booms are most suited to sites that have adequate depth of flow.  The boom and housing reduce contamination from organics by shedding debris, protecting the sensor from direct impacts by woody material, and when properly designed, they reduce hydrodynamic noise caused by turbulence and the entrainment of air or re-suspension of sediment close to the sensor.  Field personnel can retrieve the bank or bridge-mounted boom to remove debris during high flows.  The boom controls the depth of the turbidity probe and sampler intake in the stream to maintain their position above bedload transport and below the water surface.

c. Definitions

TTS - Turbidity Thresholds Sampling

SSC – suspended sediment concentration

SSL – suspended sediment load

d. Health and Safety

See the safety information in SOP 02 and 03 for safety issues when wading the stream channel.  Field personnel adding HCL to sample bottles in the field are required to ware protective clothing including safety glasses and gloves.  Field personnel lifting or carrying heavy equipment, such as sample bottles or batteries, should follow properly lifting techniques that include keeping their back straight and lifting with their legs and not twisting while lifting.

e. Cautions

Loss of data in the field can occur through carelessness.  Electronic files must be backed-up on secondary media in addition the laptop’s hard drive.  Samples may be lost, by spilling, or improperly labeled when the bottles are removed or placed in the pumping sampler.  The TTS data logger program must have the correct parameters to ensure proper operation; these include, but are not limited to, sensor multipliers and offsets, turbidity thresholds, and sampling depths.

f. Interferences

Inaccuracies in the data can result from incorrect file naming, incorrect labeling, or transposition of numbers, spillage, and discharged batteries.

g. Personnel Qualifications

Personnel must receive proper training in all facets of the field operation.  Training is administered by the Project Manager and the QA Officer will periodically review the competency of field personnel collecting field data.  Refer to the QAPP document for a description of the training and documentation.

h. Equipment and Supplies

The following equipment is required for a TTS station.  Additional instrumentation, such as rain gage, may also be installed:

- Campbell 10X or CR800 data logger

- Campbell LoggerNet software

- TTS data logger program

- Pressure transducer for measuring water stage

- In situ turbidity sensor

- Automatic pumping sampler

- Sampling boom

- Battery

- Laptop

- Cables and wiring harnesses

- Instrument shelter

- Staff plate

- Sample bottles

- Desiccant

i. Procedures

The detailed procedures are fully explained in the attached manual.

j. Data Records and Management

Bottles are labeled in the field according to the following table:

| Label Code Description | Example |
| --- | --- |
| Two-digit Hydrologic Year | 07, 08, 09 |
| Two-character sample type: IS– ISCO sample DI – Depth-Integrated sample | 07IS1413 07DI0049 |
| Four-digit unique sequential number for each sample type within the Hydrologic Year | 0001 – 9999 |
| TTS samples (additional label) HY (07,08,09) Station identifier (FTR, HHB, KRW, SFM) Bottle sequence (01-24) | 07KRW17 |

Electronic files from the data logger are maintained, with copies, on the field laptops, and office computers.  The Project Manager is responsible for maintaining the electronic files and paper field notes.  All data are filed by hydrologic year, and by station.

2. Quality Control and Assurances

QA procedures are covered in the TTS Field Manual and in the QAPP document.

3. References and Attachments

Lewis, Jack, and Rand Eads. 2008. Implementation guide for turbidity threshold sampling: principles, procedures, and analysis. Gen. Tech. Rep. PSW-GTR-212. Arcata, CA: U.S. Department of Agriculture, Forest Service, Pacific Southwest Research Station. Unofficial pre-publication layout.

Lewis, Jack. 2003. Turbidity-controlled sampling for suspended sediment load estimation. In: Bogen, J. Tharan Fergus and Des Walling (eds.), Erosion and Sediment Transport Measurement in Rivers: Technological and Methodological Advances (Proc. Oslo Workshop, 19-20 June 2002). IAHS Publ. 283: 13-20. M [294 KB]

Eads, Rand and Jack Lewis. 2003. Turbidity Threshold sampling in watershed research. In: Renard, Kenneth G.; McElroy, Stephen A.; Gburek, William J.; Canfield, H. Evan; Scott, Russell L., eds. First Interagency Conference on Research in the Watersheds, October 27-30, 2003. U.S. Department of Agriculture, Agricultural Research Service; 567-571.

Eads, Rand, and Jack Lewis. 2002. Continuous turbidity monitoring in streams of northwestern California. In: Turbidity and other sediment surrogates workshop (ed. by G.D. Glysson & J.R. Gray). 30 April - 02 May 2002, Reno, Nevada. 3 p.

Lewis, Jack. 2002. Estimation of suspended sediment flux in streams using continuous turbidity and flow data coupled with laboratory concentrations. In: Turbidity and other sediment surrogates workshop (ed. by G.D. Glysson & J.R. Gray). 30 April - 02 May 2002, Reno, Nevada. 3 p.

Lewis, Jack, and Rand Eads. 2001. Turbidity threshold sampling for suspended sediment load estimation. In: Proceedings, 7th Federal Interagency Sedimentation Conference, 25-29 Mar 2001, Reno, Nevada. [1824 KB]

Lewis, Jack, and Rand Eads. 1998. Automatic real-time control of suspended sediment based upon high frequency in situ measurements of nephelometric turbidity. In: Gray, John, and Larry Schmidt (Organizers). Proceedings of the Federal Interagency Workshop on Sediment Technology for the 21st Century, February 17-20, 1998, St. Petersburg, FL.

Lewis, Jack. 1996. Turbidity-controlled suspended sediment sampling for runoff-event load estimation. Water Resources Research 32(7): 2299-2310.

Lewis, Jack, and Rand Eads. 1996. Turbidity-controlled suspended sediment sampling. Watershed Management Council Newsletter 6(4): 1&4-5.

Turbidity Threshold Sampling

Field Manual 5.X

LoggerNet Software 3.X

Campbell CR10X Data Logger

© 2007 RiverMetrics LLC

rivermetrics@gmail.com

www.rivermetrics.com

(503) 435-7516

### Revised: 08/26/2007

- Table of Contents

Observing and recording station information	3

Connecting to the data logger and starting LoggerNet	3

Establishing communications between the laptop and data logger	3

Viewing, observing, and recording information	4

Determining what needs to be done	6

Collecting an AUX sample	7

Collecting a DI sample	7

Dumping the data	9

Changing the ISCO pumping sampler bottles	11

Maintaining the station	12

Plotting the data	13

Complete the field form	14

Backing up the data file	15

Changing values in the Numeric window	15

Loading a new program revision	16

Viewing the data file	17

Basic Description of the TTS Program	17

Troubleshooting	20

Programming the ISCO 6712	23

LoggerNet station setup	24

- Observing and recording station information

  - On the field form record the Date, Time, Observer initials.

  - Note the presence of sediment and/or debris that is in the channel or obstructing the turbidity probe (to be cleared later) and read the stage, but only record within 5 minutes of a data logger wakeup.

  - Observe current ISCO pumping sampler display and check the samples.  Do not disturb the ISCO pumping sampler within +/-2 minutes a wakeup. If sample volumes are too high or too low, go to Troubleshooting.

  - Prepare to change the ISCO pumping sampler bottles if needed (instructions in Section 9) and perform routine station maintenance.

- Connecting to the data logger and starting LoggerNet

  - Connect the communications cable and SC32B serial interface between the laptop and the data logger.

  - Connect the mouse (optional).

  - Open the laptop and turn it on.

  - Click on the LoggerNet icon:

- Establishing communications between the laptop and data logger

  - Click on the Connect icon.

  - Select the Station Name from the list and click on Connect:

  - Make sure the laptop’s clock is accurately set to the telephone time recording before setting the data logger’s clock.  Compare the data logger clock to the laptop clock. If the error is greater than 30 seconds, adjust the clock after you transfer the data.

- Viewing, observing, and recording information

  - Click on the Numeric 1 button:

The Numeric window displays the values from the most recent wake-up.

  - Record the current time and information from the following fields in the station Numeric window, onto the field form:

- bat_volt (battery voltage): if voltage is less than 12.0, notify the project supervisor.

- stage (electronic stage): Compare this value with the observed stage from the staff plate reading. If the stage differs by more than 0.02 feet, make note, view plot, and take further action as necessary, after collecting the data.  Be cautious about adjusting the offset during high flows.

- med_turb (turbidity): The reading is valid when stage is above the value set in wipe_stg.

- nxt_bot (next Campbell sample number): Compare this value with the ISCO pumping sampler display. If numbers differ, go to Troubleshooting after collecting the data.

- dump_cnt (data dump number): Compare this value with the dump number from the previous field form. If the number displayed on the data logger is incorrect, change it now.

- wtemp (water temperature).  The DTS-12 must be fully submerged for this temperature to be valid.

- rain (rainfall during previous interval).  One tip is equal to 0.01” of rainfall.

- min_stg (minimum stage).  Both the turbidity sensor and ISCO pumping sampler intake are adequately submerged above this stage.  To modify the value double-right click on the current value, then type the new value and press [Enter].

- stg_off (stage offset).  Enter a value to make the electronic stage agree with the staff plate.  Avoid making changes during high flows or when it is difficult to obtain an accurate staff plate reading.

- wipe_stg (wiper activation stage).  This controls the wiper independently from min_stg.  Set the stage value so the wiper is deactivated just before it becomes dry.  For example, min_stg = 0.52 and wipe_stg= 0.39.  No samples would be collected below 0.52 but the wiper would continue to work until the stage dropped to 0.39.  Wipe_stg should always be ≤ min_stg.  This prevents the optics from fouling below min_stg if conditions permit.

- Determining what needs to be done

  - Decide if a depth-integrated (DI) manual sample is required using the following criteria:

- The current turbidity is within a range that has not been sampled and there is adequate sampling depth.

If a DI sample is required, go to Section 7.0. If not, go to the next step.

  - Decide if an auxiliary (AUX) sample is required using the following criteria:

- If the next sample number displayed on the ISCO pumping sampler does not agree with the number in the Numeric window, or

- The ISCO sample volumes are inadequate, or

- The stage is receding towards the minimum stage and there are less than 4 samples from the current storm event.

If an AUX sample is required, go to Section 6.0. If not, go to the next step.

  - Decide if a data dump is required using the following criteria:

- More than two weeks have passed since the last dump, or

- More than 14 samples have been collected and additional sampling is expected before the next site visit, or

- You are going to load a new program revision, or

- The site is experiencing equipment problems.

If a Dump is required, go to Section 8.

- Collecting an AUX sample

  - Open the Ports And Flags window:

  - In the Ports And Flags window, click on the button to the left of AUX (the button will turn from black to green):

NOTE: You can click again on the button to deselect it if you made a mistake. Do not select or deselect any other buttons.

- The program collects the AUX sample within one minute after the next wakeup, and automatically turns the AUX button from green to black.

  - After the AUX sample is collected note the value of “nxt_bot” in the Numeric window, confirm that it agrees with the ISCO “next bottle”, and record the value in the station notebook.

  - Decide if a Data Dump is needed.

- Collecting a DI sample

  - Prepare the DI bottle label with station letters, data dump number, matching ISCO pumping sampler bottle number, and time of pumped sample collection (next wake-up).

  - Open the Ports And Flags window:

  - In the Ports And Flags window, click on the button next to DI to select  the DI flag (it will turn from black to green):

NOTE: You can click again on the button to deselect it if you made a mistake. Do not select or deselect any other buttons.

- The program will collect the companion pumped sample within one minute after the next wakeup, and will automatically turn the DI button from green to black.

  - Place the bottle in the depth-integrated sampler. Initiate depth-integrated sampling (DI) as soon as the ISCO pumping sampler stops sampling if the sampling location is upstream of the ISCO pumping sampler intake, or time the mid-point of the depth-integrated collection to coincide with the ISCO pumping sampler collection when the DI sampling location is below the intake.

  - After the DI and ISCO pumping sampler samples have been collected note the value of “nxt_bot” in the Numeric window, confirm that it agrees with the ISCO “next bottle”, and then record it in the station notebook.

  - Decide if a Data Dump is needed.

- Dumping the data

  - If a wake up has occurred since your initial observation of the Numeric window, re-examine the values in the Numeric window AND note the “next bottle” number on the ISCO display.  Update any changes to the station notebook and note the current time.

  - A sample may be collected at the next wakeup.  Be sure to allow yourself a full wakeup interval to service the ISCO pumping sampler and bring it back online before that next wakeup occurs.

  - Open the Ports And Flags window:

  - Be sure that the next wakeup is at least 3 minutes away, then click on the button to the left of DUMP (the button will turn from black to green):

- The dump_cnt value displayed in the Numeric window will increment by one at the next wake up.

NOTE: You can click again on the button to deselect it if you made a mistake. Do not select or deselect any other buttons.

  - Change ISCO pumping sampler bottles and perform the station maintenance tasks, noting times and details in station notebook.

  - Click on the Data Collection button labeled Custom.

  - Edit the date within filename to reflect the date the file started—leave the remainder of the path as it is displayed.  Always use the directory structure: C\:newdata\ssyymmdd (where ss are two letters of the station, yy is the year, mm is the month, and dd is the day, chose .dat as the extension).

CAUTION!  If this station was dumped earlier on today’s date, this naming convention will result in duplicate filenames!

Instead, the new file name is:  ssyyddnn.dat (where: yy is hydrologic year and nn is the dump_cnt)

- Follow the file transfer in the % Collected bar until it displays 100% and disappears from the screen:

- When file transfer is complete, the program will automatically turn the Dump button from green to black at the next wake up.

- If you need to correct the data logger clock, do so now after verifying that the laptop time is correct and that the time discrepancy does not straddle a wakeup (for example: 10:18:30 = data logger clock and 10:20:30 = laptop clock.)  Click on Set Datalogger Clock.  See Section 20 to understand how to use the clock time zone offset to account for differences between standard and daylight savings time.

- If you need to change min_stg, (minimum stage) or stg_off (stage offset), do so now by following the directions in that section.

- Close the Ports / Flags window by clicking on the X in the upper right corner on the window.

- Click on the Disconnect button:

- Close the LoggerNet window by clicking on the X in the upper right corner.

- Disconnect the communications cable from the data logger.

- Go to Section 11.

- Changing the ISCO pumping sampler bottles

- Prepare replacement sample bottles (this step should be done before starting the data dump):

    - Label each bottle “sss – nn”, where

- “sss” is the three-letter gaging station ID, and

- “nn” is the sequential bottle number (01 to 24)

- 9.1.2.	Remove the caps.

- Unfasten the bale latches on the sampler mid-section, and lift it off the sampler base.

- Tightly cap each bottle containing a sample before remove the bottle from the base.

- Place the sample bottles on the designated level shelf.

- Insert the empty, uncapped replacement bottles in the correct sequence, matching the bottle number on the tape label to the number on the sampler base.

- Re-assemble the ISCO pumping sampler.

- Reset the sampler to bottle position 1:

Press red STOP button

Press [↑] to scroll up until the “RUN PROGRAM” is blinking

Press the ENTER key

The ISCO pumping sampler will display “BOTTLE 1 AFTER 1 PULSE”

- While holding each sample bottle on the sloping shoulder, loosen and re-tighten the cap (to achieve an accurate volume).

- Mark the sample volume level by placing a piece of tape on the wider side of the bottle and carefully marking the bottom of the water meniscus.

- Write the data dump number, dd, on each label in front of the station letters.

- On bottle 01, within parentheses, write the number of samples (x) that where removed.

Label Format for bottle 1: dd-ss-nn (x)

- Maintaining the station

- Perform the following station maintenance tasks:

  - Check the turbidity probe and sampling boom between wake-ups:

- Removing debris from housing and leading edge of the boom, and

- Remove sediment from inside housing by rapidly pushing the boom up and down stream, and

- Adjust the boom height, if necessary, but use caution since there is a tendency to over-correct the height resulting in a setting that is too high or too low.  Most sites do not require adjustment.

- During non-storm visits, clean the optics if necessary

- Check the measurement section between wakeups:

- Remove branches and debris, and

- Read and record the staff plate a second time after clearing measurement section

  - Check/replace station batteries as needed by following these steps:

- If the 12V battery is <12.0 volts only replace the battery while another power source is connected (9V or 12V battery, or solar panel during the day).

- Turn off the ISCO pumping sampler if a 9V back-up battery is in use.

- If a 9V battery is used:  Set the replacement battery next to the battery being removed and quickly remove the power leads and reconnect to the replacement battery.

- If only 12V batteries are used:  Set the replacement 12V battery next to the discharged 12V battery and connect them in parallel, then disconnect the discharged battery.

- Turn on the ISCO pumping sampler if a 9V battery is in use.

- Check the data dump number to make sure it is correct, if not edit the value in the numeric window.  If the dump number reverts to “1” after the first wake-up it means that power was lost.

  - Replace desiccants when the middle circle on the moisture card inside the enclosure turns pink.

  - Adjust, troubleshoot, or replace equipment as needed.

- Plotting the data

- Double click on the icon R icon.

- At the File Selection? prompt, type in the number corresponding to the current filename.

- Wait until the complete TTS window appears.  Sometimes it is necessary to drag the dialog window with the mouse to get the window to fill completely.

- Click on the Data Type button and choose a plot type (usually stage-turbidity).  The plot parameters will then auto-fill:

    - Verify that the End Date/Time agrees with today’s date and the time of the wake-up prior to today’s data collection,

CAUTION! There should be not be unexplained gaps in the data

- Click on Plot and wait for the plot to appear...

- Click on OK at the Legend prompt and, if plotting more than one variable, click on the plot where you want to place the upper left-hand corner of the legend.

- Examine the plot to see if the data appears to be correct. Edit date/time and axes as needed in the TTS window for better resolution and click again on Plot to evaluate suspicious data.

NOTE:  The TTS window may disappear to the bar at the bottom of your screen.  Just click on the red TTS to restore it.

- When finished viewing the data, click on Quit in the TTS window.

You should now have a file selection prompt in the R console

- Answer NO at the SAVE WORKSPACE IMAGE popup

- Go to the next section.

- Complete the field form

- Fill out or strike-through ALL fields on the field form:

- Use the time, stage readings, and bottle values that were made after any AUX or DI sample collection.

- If a data dump was done, staff plate readings should be made within 5 minutes of the final record of the data dump and the time on the field form should be recorded as the time of that final record.

- Where the left, right, or inside staff plate is nonexistent, enter NA

- Enter NA for the Spare 12v Battery Voltage field if there is no spare 12v battery.

- Use the appropriate letter designation for the bottle status of each of the 24 sample bottles.

- Make notes on the bottom of the form.  Be thorough and make concise comments on what was done and observed during this visit.  Use the back of the form if you need more space.  Note any conditions that may affect the data quality.  Make a list of supplies needed at the next site visit.

- Start a new field form for the next site visit.  Fill out the following fields below and complete other header fields if equipment parameters values have changed, or if it is the first field form of a new hydro year:

    - Site (xxx): first three letters of station name

    - Hydro Year (yyyy)

    - Data Dump (nn)

- Backing up the data file

  - Insert a memory stick or flash card in the laptop

  - Copy the file from C:\newdata onto the device.

  - Disconnect the device and remove it from the laptop.

- Changing values in the Numeric window

Explanation:  Program parameters in the Numeric window may be changed to temporarily update the TTS program on the data logger.

CAUTION:  Changing a value in the Numeric window modifies the program operation but not change the program in the data logger, laptop, memory sticks, etc.  If the TTS program is reloaded, the changes made in the Numeric window will be lost.

- Place the mouse pointer in the field you wish to change and double-click on the value

- Press the backspace key to erase the old value

- Type the new value.

- Press [ENTER] for the new value to take effect.

- Exit the Numeric window and LoggerNet software now, or wait until the next wakeup to see that your change has taken effect.

- Loading a new program revision

Make sure that the TTS data file has been successfully offloaded to the laptop and backed-up on a memory device before continuing (the file in the data logger will be lost if the length of the program is changed) .

Verify that the current values for min_stg, stg_mult, and stg_off are correctly recorded on the station note sheet.

- Select the appropriate program “sss_n_n.dld” (where sss are the three station letters, n is the revision numbers, and .dld is the compiled version of the program) from the list and click OK.

- Click on the Send button located on the Connect screen.

- When the warning window appears with the message about data loss, click on OK.  (You should have already dumped the data and backed it up).

- The .dld file will now be sent to the data logger:

NOTE:  The first wake-up after the program is loaded will usually be skipped.  For example, if you load the program at 1006 the first wakeup will likely occur at 1020.

- Wait until the Numeric window displays the values generated at the first wakeup.

- You may need to update the values in the Numeric window to match what is posted in the station notebook.  You will need to revise the data dump number (always reverts to 1) unless it is the first dump of the year.

- Viewing the data file

- Navigate to the C:\newdata directory.

- Open the current file by right clicking on it and choosing “Send To Notepad”.

Here is the data format:

Site ID,Year,Julian day,Time,Dump Number,Bottle Number,Threshold Code, Sample Code,Stage,Turbidity,Water Temp,Air Temp,Rain (where Water Temp, Air Temp, and Rain are optional fields)

- Click on Exit.

- Choose No to Save Changes prompt.

- Basic description of the TTS program

What the TTS Program Does:

At each wakeup:

- The DTS-12 returns several statistics to the data logger.  The program retains only the median turbidity value and then adds turb_off to the value.  This allows the user to compensate for sensor drift by adjusting turb_off.

- Stage is measured with a pressure transducer.  Raw voltage is converted to stage in feet using the program parameters stg_mult and stg_off (both are sensor/site-specific values).

- The program may trigger a pumped sample if the threshold sampling criteria (stage and turbidity) are satisfied or an AUX or DI sample has been selected by the field crew.  If the ISCO pumping sampler is set to sample, it will do so 35-40 seconds after the wakeup, 50 seconds with the DTS-12 connected.

- The data record is written to memory.

- The data logger goes to low power mode until the next wake-up.

Definitions of Sampling-related Program Variables:

thr  = Turbidity threshold for sampling

thr_count  = Threshold count

0	Sample previously collected, or startup condition

1	First interval above next threshold

2	Second interval above next threshold, sample is collected

rev_count = Reversal count

0	Previous reversal, or startup condition

1	First interval after a change in turbidity direction

- Program registers reversal of turbidity direction if change is great enough.  Sample is collected if threshold was crossed since previous max (on switch to falling) or min (on switch to rising).

thr_code = Threshold code

0	Baseflow (stage below minimum stage)

1	Rising turbidity

2	Falling turbidity

3	Unknown turbidity, not yet defined as rising or falling

smp_code = Sampling code

0	No sample collected

1	Threshold sample

2	Depth-integrated sample (DI)

3	Auxiliary sample (AUX)

4	Start-up sample

5	Overflow sample, turbidity exceeds turbidity probe maximum,

(sampling interval is user selectable)

- Sampling Conditions (Modes):

- Baseflow:  This condition occurs when the stage is less than the minimum stage (min_stg).  Minimum stage is defined as the lowest stage where both the ISCO pumping sampler intake and turbidity probe are submerged and functional.  No threshold sampling takes place in this mode.  The variable thr_code = 0.

- Rising:  Turbidity Mode automatically becomes Rising at first interval above baseflow.  If turbidity is also above the first threshold and no rising thresholds have been sampled in the past 3 hours, a start-up sample is collected.  For subsequent rising turbidity mode samples, current turbidity must equal or exceed the next rising threshold for 2 intervals.  The variable thr_code = 1.

- Reversals:  Turbidity Mode switches between Rising and Falling.  The turbidity must change direction for at least 2 intervals, AND drop 10% from the prior peak or rise 20% from the prior trough (user selectable), but at least 5 NTUs in either case (user selectable).  A sample is collected if a threshold has been crossed since the previous peak or trough.

- Falling:  A falling turbidity mode sample is collected when the current turbidity drops below the next falling threshold for 2 intervals.  The variable thr_code = 2.

- Repeat samples:  When conditions for a threshold or reversal sample are otherwise met, a sample will not be collected if the threshold that was crossed has already been sampled in the past 18 (rep_wait) intervals (user selectable).

- Overflow:  Occurs when the turbidity sensor’s output range is exceeded by the value held in the variable (ntu_limit for the DTS-12 and mv_limit for the OBS-3).  In this mode, (lim_skip) intervals will be skipped between each fixed time sample interval.

- Unknown:  Turbidity trend not yet defined after a cold start.  The variable thr_code = 3.

- TTS Sampling Thresholds

Standard thresholds are listed below.  The number of thresholds and their values may vary depending on site conditions and may change during the season.

Rising Thresholds       Falling Thresholds

| 20 | 1129 |
| --- | --- |
| 59 | 1014 |
| 120 | 906 |
| 201 | 802 |
| 303 | 706 |
| 427 | 616 |
| 571 | 532 |
| 736 | 454 |
| 922 | 382 |
| 1129 | 316 |
|  | 2571 |
|  | 204 |
|  | 156 |
|  | 116 |
|  | 81 |
|  | 52 |
|  | 30 |

- Troubleshooting

If possible, retrieve the data before attempting to troubleshoot.

Record everything: symptoms, actions, and results.

Identify, repair, and replace malfunctioning equipment.

- Diagnostics Programs

Note:  The following diagnostic programs enable repeated tests of a data logger function over a short time frame:

- DRUCK5PSI.csi – Reports mean stage from pressure transducer at 10-second intervals.

- DTS12test.csi – records DTS-12 turbidity statistics at 60-second intervals using C2

- ISCO pumping sampler_c1.csi – Tests ISCO pumping sampler connected to Port 1 (C1), runs at 15-second intervals

- Problem:  Data logger not connecting or dropping connection

    - Check cables between data logger and laptop.  Be sure to check ALL connections.

    - Check Battery voltages with voltmeter.  If 9V is in use it must be at least 9.0 volts.  Replace any low batteries.

    - Problem:  Station Numeric window does not contain necessary variables or values

    - If the variable names are present but the values are blank, you have not successfully connected to the station data logger (see above).

    - If the variable names in the Numeric window = “Inloc_1”, “Inloc_2”, etc. then the laptop is not setup correctly for this station.

- Confirm that you have selected the correct program file name.

- At the bottom of the Numeric window, click on ADD button.  A second window will open on the left showing all the variables.  This window is titled “InLoc list”.  Scroll down the list and drag and drop the designated variables (dump_cnt, stage, and so on) into the Numeric window, see example in Section 4. The values for each variable should appear as they are dropped on the Numeric window.

- Close the “Inloc List” window.

    - If the variable names appear, but the values are equal to zero, the data logger is in a “Cold Start” Mode.  The values should fill at the second wakeup.  Wait and see.  You may have to update some parameters (for example, stg_off.) if they were not current in the program.

    - Problem:  Dump_cnt is incorrect

Note:  The dump_cnt reverts to 1 at a “cold start” (after loading a new program revision or after a power loss).  To correct the dump_cnt See changing values in the Numeric window, Section 15.  The station notebook and field forms track dump numbers.

- Problem:  Displayed stage is not correct within acceptable limits

NOTE:  During the first storm flow of the season, the electronic (Campbell) stage is frequently off and the transducer index (stg_off) will need to be adjusted.  If a stage error is noted after an initial seasonal adjustment has been done, additional troubleshooting should be done.

    - To adjust the stg_off in the Numeric window:

- Do a data dump first.

- Re-open the Numeric window

- Offset correction = (staff plate) – (stage)

- Add this correction to the current value of the stg_off

- Wait for the next wake-up and confirm that the electronic stage agrees with your current observed stage.

- Be sure to record the change on the station notebook and on the field form.

    - Check the previous field forms and station notebook to determine if changes were made to the stg_off.  The offset may have been set incorrectly or reverted to a previous value.

    - Verify that the transducer is secured.  If you suspect it has moved, re-secure the transducer and reset offset.

    - Inspect the transducer cable for damage.

    - If the stage plot indicates drifting or erratic stage over time suspect a failing pressure transducer.  Replace the pressure transducer.

    - If the stage error occurs after a high sediment event, check for debris and flush the intake if clogged.  Note action, extent of obstruction, and post-cleaning stage on Field Form.

    - If the transducer has been replaced and the problem persists, suspect a data logger problem.

    - Problem:  ISCO pumping sampler bottle volumes too low (empty) or too high (overflow)

Volumes may be low when ISCO pumping sampler intake is marginally submerged.  Avoid calibrating sampler at stages very near to the station minimum stage.

Ideal sample volume is 1/3 of bottle capacity (~330ml).  Minimum is 1” depth.

- Check station status sheet and ISCO pumping sampler controller for proper settings.

- Inspect for cracked and leaking ISCO pumping sampler bottles.

- Check intake tubing for leaks or kinks.  Make sure connections are tight.

- Check inside ISCO pumping sampler pump housing for worn pump tubing.  An air leak at the junction of the intake and pump tube will cause low volumes.  Use ONLY stainless steel hose clamps, NEVER use the plastic clamps that come with the tubing connector

- Inspect ISCO pumping sampler intake for obstruction (if stage permits).

- Over- or under-filled bottles in an ISCO pumping sampler can be caused from a liquid level sensor failure, air leaks, air bubbles from high stream velocity, or using ¼” id intake instead of 3/8” id intake.  All of the above conditions can be fixed except for air bubbles in high velocity streams.  The liquid level detector should be used for all stations so the intake line can be rinsed before sampling to reduce cross-contamination between samples.

- Enter the configuration mode and select the following:

- Disable “Liquid detector”

- Calibrate sample volume.

- Problem:  ISCO pumping sampler over-sprays or water in base of sampler when sample volumes are adequate

- Adjust distribution tube-end to extend 1/16" beyond the distributor arm.

- Inspect distributor arm for loose connections, binding, or twisted tube.

- Check for cracked and leaking bottles.

- Problem:  Data won’t plot in R

- Exit the Rplot program and retry.  Sometimes multiple copies of Rplot are opened and memory is insufficient to run the program.

- Press [alt] [tab] to move through all operating windows on the laptop.  Be sure to close all copies of Rplot, as well as the Paradox program if it is running.

- Retry by double clicking on the R icon.

- Programming the ISCO 6712

- Connect to a power source and turn on the ISCO pumping sampler.  (Refer to the ISCO pumping sampler manual for the description of the keypad.  Pressing the “Enter” key after an entry is implied in the instructions below).

- Follow the 6712 Manual steps for setting the “Clock and Calendar” (Other Functions – Maintenance- Set Clock).  See the corresponding steps in the ISCO pumping sampler manual for a more compete programming instructions.

- At the main menu, go to the “Extended Programming” screen by typing “6712.2”.

- Select the program “Extended 1”.

- Under “Program Name” you can change the program name to “TTS”.

- You can rename “Site Description” of “Factory xxx” to the three-letter station identifier.

- At the “Select Units for Length” prompt chose “ft”.

- Select a 10-minute data interval (this won’t actually be used).

- Select the bottle configuration: 24, 1000 ml BTLS (or the appropriate size if using glass bottles or 500 ml plastic wedge bottles).

- Enter the “Suction Line Length”: xx in feet (this is site specific).

- Select “Auto Suction Head”.

- Under “Rinse Cycles” select “1”.

- Under “Retry when Sampling” select “0”.

- Select “One-part program”.

- Under the Pacing option select: Flow Paced, change to “1 Pulse”.

- Select “No Sample at Start”.

- For sample distribution select: “Sequential 350 ml samples” (may be less if using small bottles).

- Under “Enable: None programmed”.

- Skip through the remaining “enable” fields.

- Select “No Delay to Start”.

- Select “Run Program Now”.

- LoggerNet station setup

- Click on the toolbar EZ button:

- You can click on the Help button at the bottom of the window at any time.  When finished with each window click on the Next button.  Click on the Add button now:

- Select the data logger model and name the station in the Station Name box.

- Select Connection Type as Direct Connect:

- Select the Communications Port and add a 2-second delay.  When using a USB port you may have to try different ports until you get a connection:

- Use the default settings but you may want to add an additional 2-second Response Time:

- The Setup Summary page reviews your choices.

- Test the communications.  If you fail to get a connection use the Previous button to go back the COM Port Select and select a different port:

- Set the data logger clock.  Using the Time Zone Offset feature will allow the laptop’s clock to be set according the time recording on the phone.  The offset will then match the data logger and laptop clock.  In this example, the clock is offset (-1) for Daylight Saving time.  When daylight savings is not in effect set the offset to “0”:

- Load the program into the data logger.  Use the Select and Send Program button to browse to the correct directory in the laptop where the station’s program is located:

- Data file storage area.  Storage Area 1 is selected as the default.  Under Output Filename change the directory to: C:\newdata

- Click on the Finish button.