Surface Water Ambient Monitoring Program (SWAMP)

# Suspended Sediment and Turbidity Monitoring

# Freshwater Creek and Elk River

STANDARD OPERATING PROCEDURE 04

FIELD INSTRUMENTATION

Prepared for:

Salmon Forever

Arcata, CA

Prepared by:

Rand Eads

Surface Water Specialist

RiverMetrics LLC

PO Box 507

Lafayette, OR 97127

Version 1.0

February 2008

1. Procedures

a. Scope

This document covers the installation, maintenance, and troubleshooting of field instrumentation that includes pressure transducers, turbidity sensors, pumping samplers, and rain gages.

b. Method Summary

c. Definitions

SDI-12 protocol: digital data protocol used by the DTS-12 turbidity sensor

Control port: a data logger port that activates the pumping sampler

Tipping bucket: the device in the rain gage that measures volume

Pulse count: one pulse count from the rain gage is 0.01” of rain

Single-ended sensor: OBS-3 turbidity sensor, analog connection

Automatic suction head: a setting in the pumping sampler to control volume

Vent tube: vents pressure transducer to the atmosphere pressure

d. Health and Safety

There are no health and safety concerns with these protocols.  Turbidity sensors, mounted inside the sampling boom can be lifted mechanically so field personnel can service the boom without entering the stream.  The pressure transducer can be removed from the conduit for servicing, usually from the bank without entering the stream.

e. Cautions

Data can be compromised from improper placement of the sensor, failure to secure the sensor, or an incorrect calibration.  Additional data problems can occur if the sensor is connected to the data logger incorrectly or the wrong sensor parameters are used in the data logger program.

f. Interferences

Inaccurate calibration of the sensor or incorrect ISCO sample volume can degrade data quality.

g. Personnel Qualifications

Personnel must receive proper training in all facets of the field operation.  Training is administered by the Project Manager and the QA Officer will periodically review the competency of field personnel calibrating or installing sensors.  Refer to the QAPP document for a description of the training and documentation.

h. Equipment and Supplies

The following equipment is in use:

- Druck 1830 pressure transducer

- ISCO 3700 or 6712 automatic pumping sampler

- Campbell TE-525 tipping bucket rain gage (6” and 8”)

- D & A OBS-3 or FTS DTS-12 turbidity sensor

- Additional ISCO sample bottles

- Desiccant

i. Procedures

Druck 1830 pressure transducer

Calibration

The Druck pressure transducer should be check annually for drift in the slope (multiplier) and stability of the output.  A cloth measuring tape or marked dipstick is used to position the transducer in a 1” tube filled with water.  The tube is held vertically and the first measurement is taken at 0.3 feet.  Subsequent measurements are taken at 1, 2,3,4, and 5 feet.  The pressure transducer is lowered to depth that is read from the bottom of the meniscus.  The third reading (mV) at each depth is recorded.  The data, output vs. depth, is regressed in an Excel spreadsheet.  Fit the data with a linear trend line and display the equation and r2.  The fit should be linear and the r2.should approach 1.  Any outliers should receive a repeat measurement.  The intercept (stage offset) is ignored from the results because it is set in the field after reading the staff plate.

Maintenance

Install fresh desiccant in the enclosure every 2 to 3 weeks to prevent moisture from entering the vent tube and causing an electronic failure.

Installation/Mounting

The pressure transducer should be mounted in conduit and close to the staff plate.  The sensor should be mounted so the lowest stages can be measured, yet the sensor should not be covered by sediment.  The opening of the conduit should be at right angles to the flow.

Campbell TE 525 tipping bucket rain gage

Calibration

The tipping bucket rain gages should be checked for annually for accuracy by dripping a known volume of water, at a specified rate, through the gage.  See the manual for the recommended drip rate and volume.  The numbers of tips are recorded in the data logger using the “raintest” program (see reference below).

Maintenance

Remove debris and dirt from the screen, funnel, and tipping mechanism every two to three weeks.  Maintain a 2:1 horizontal distance to height ratio for nearby obstacles (trees and shrubs).

Installation/Mounting

Mount the rain gage in a clearing approximately 4 feet above the ground.  The top of the funnel must be level.  The rain gage must not shake in the wind.

ISCO 3700 and 6700 pumping sampler

Calibration

The sample is volume is set in the ISCO controller program.  A minimum of 125 ml is required, and a nominal volume of 350 ml is desired.

Maintenance

Check the intake tube periodically for sediment or algae accumulation.  Change the desiccant in the controller annually.  Replace the pump tubing when the controller displays a warning (also reset the warning counter).

Installation/Mounting

Maintain a head height less than 25 feet.  The intake line must be routed in conduit and the sampling boom to eliminate dips in the line.  The intake line must run continuously downhill to the stream.  Use 3/8” id tubing from ISCO.  Use the Auto Suction Head option unless velocities create excessive air bubbles.  Set the Rinse to 1 and Retries to 0.

OBS-3 turbidity sensor and DTS-12 turbidity sensor

Calibration

Calibration should be performed annually by the manufacturer.

Maintenance

Remove algae or macroinvertebrates from the optical surface during each site visit with a soft toothbrush and non-abrasive liquid soap.  Remove debris from the sampling boom.  Insure the sampling boom is located in the optimal stream location and depth.  Replace the wiper on the DTS-12 annually or if it shows wear.

Installation/Mounting

Mount the sensor inside of a protective housing secured to a sampling boom.  The sensor should be located in the thalweg and approximately at mid-depth.

j. Data Records and Management

Calibration data is recorded in the station notebook, calibration forms, or on a spreadsheet.  Calibration data is maintained by the Project Manager.  Pressure transducer multipliers are derived from regression analysis using Excel spreadsheet.

2. Quality Control and Assurances

During the pressure transducer calibration allow the electronic signal to settle (for 3 readings) before recording sensor output at each depth.  The r2 should be no less than 0.99 for a properly functioning sensor.  If one or more depth measurements do not fit the trend line repeat the measurements at those depths.  If the data are not within the constraints listed above consider the sensor unstable.  The sensor may be left at a fixed depth in the calibration tube for a day or more to check if the output is stable by plotting the series data.

Use three replicates when delivering a know volume of water to the tipping bucket rain gage.  If the tips recorded indicate, an error larger than the specifications the stop screws may be adjusted to increase or reduce the volume per tip.

3. References and Attachments

http://www.campbellsci.com/documents/manuals/te525.pdf

http://www.campbellsci.com/documents/manuals/te525.pdf

http://www.campbellsci.com/documents/manuals/obs-3+.pdf

http://www.ftsinc.com/

http://www.isco.com/pcfiles/PartPDF3/UP000TBF.pdf

http://www.isco.com/pcfiles/PartPDF4/UP000ZOT.pdf

http://www.fs.fed.us/psw/topics/water/tts/software/dts12test_csi.txt

http://www.fs.fed.us/psw/topics/water/tts/software/isco_c1_csi.txt

http://www.fs.fed.us/psw/topics/water/tts/software/raintest_csi.txt

http://www.fs.fed.us/psw/topics/water/tts/software/druckcal_csi.txt