Surface Water Ambient Monitoring Program (SWAMP)

# Suspended Sediment and Turbidity Monitoring

# Freshwater Creek and Elk River

Quality Assurance Project Plan

Prepared for:

Salmon Forever

Eureka, CA

Prepared by:

Rand Eads

Surface Water Specialist

RiverMetrics LLC

PO Box 507

Lafayette, OR 97127

Final Version 1.0

April 15, 2008

# Group A: Project Management

# 1.0 Approval Signatures

________________________________________________          ___________________

Jesse Noell, Project Director				                                 Date

Salmon Forever

________________________________________________          ___________________

William Ray, State Board Quality Assurance Program Manager                Date

SWRCB, Division of Water Quality, Sacramento, CA

________________________________________________          ___________________

Tyler Ledwith, QA Officer					                       Date

Natural Resources Services, Redwood Community Action Agency

________________________________________________          ___________________

Don Allan, Co-Director							Date

Natural Resources Services, Redwood Community Action Agency

________________________________________________          ___________________

Adona White, Water Resource Control Engineer				Date

North Coast Region Water Quality Control Board

________________________________________________          ___________________

Kristi Wrigley, Laboratory Director					          Date

Salmon Forever

# 2.0 Table of Contents

Suspended Sediment and Turbidity Monitoring	1

Freshwater Creek and Elk River	1

Group A: Project Management	2

1.0 Approval Signatures	2

2.0 Table of Contents	3

3.0 Distribution List	5

4.0 Project/Task Organization	6

5.0 Study Areas and Problem Statement	9

6.0 Project Description and Schedule	10

7.0 Quality Objectives and Criteria for Measurement Data	13

8.0 Special Training Needs and Certification	14

9.0 Documents and Records	15

Group B:  Data Generation and Acquisition	16

10.0 Sampling Process Design	16

11.0 Sampling Methods	16

12.0 Sample Handling and Custody	17

13.0 Analytical Methods	18

14.0 Quality Control	18

15.0 Instrument Testing, Inspection and Maintenance Requirements	20

16.0 Instrument/Equipment Calibration and Frequency	21

17.0 Inspection/Acceptance of supplies and Consumables	22

18.0 Non-Direct Measurements	22

19.0 Data Management	23

Group C: Assessment and Oversight	23

20.0 Assessment and Response Actions	23

21.0 Reports to Management	25

Group D: Data Validation and Usability	25

22.0 Data Review, Validation and Verification Requirements	25

23.0 Validation and Verification Methods	26

24.0 Reconciliation with Data Quality Objectives	26

References	27

Acknowledgements	29

Appendix	30

# 3.0 Distribution List

| Title | Name | Phone No. | Hard Copies |
| --- | --- | --- | --- |
| Project Director | Jesse Noell | 707-445-9555 | 1 |
| Laboratory Director | Kristi Wrigley | 707-443-1496 | 1 |
| QA Officer | Tyler Ledwith | 707-269-2058 | 1 |
| Grantee Grant Manager; Co-Director RCAA | Don Allan | 707-269-2063 | 1 |
| Water Resource Control Engineer | Adona White | 707-479-2342 | 1 |
| State Board Quality Assurance Program Manager | William Ray | 916-341-5583 | 1 |

# 4.0 Project/Task Organization

4.1 Involved Parties and Roles

Redwood Community Action Agency (RCAA) is administering the grant received from State Water Resources Control Board, and the agency has entered into a contract with Salmon Forever to perform the project work.  Don Allen, Co-Director Natural Resources Services, RCAA, is the principal contact.

## Contractor

Salmon Forever is the primary contractor.  Salmon Forever is responsible for the collection and analysis of data from six gaging stations:  two on Freshwater Creek and four on Elk River.  Salmon Forever is also responsible for purchasing equipment, training staff and participants, processing and analyzing samples, producing reports, hiring subcontractors for specialized work.

### Subcontractors

Subcontractors provide specialized skills such as equipment fabrication, installations, training, preparing manuals, surveying, gaging station maintenance, discharge measurements, and data analysis.  Subcontractors work directly for Salmon Forever.  Salmon Forever provides the specifications for each task, timelines, and reviews the work for completeness.

Jesse Noell, Project Director

The Project Director is responsible for question formulation, parameter selection, and developing the sampling design with oversight of the Technical & Scientific Advisory Panel.  The Project Director reviews field and laboratory data for QAPP objectives.  The Project Director is responsible for writing and distributing reports.  The reports will be reviewed by the Technical & Scientific Advisory Panel, RCAA, and North Coast Regional Water Quality Board (NCRWQCB).  Recommendations from these groups will be used to modify the project goals, guidelines, and data collection methodologies.

The Project Director supervises the Station Masters who are responsible for the operation of the gaging stations.  In addition, the Project Director provides field and safety training, on-call technical support; checks field samples; verifies field data, and transports field samples to the laboratory.  He is responsible for the chain-of-custody for all field samples.  The Project Director supervises subcontractors performing fieldwork.

Kristi Wrigley, Laboratory Director

The Laboratory Director supervises and oversees all laboratory personnel including the Laboratory Manager, students, and volunteers.  The Laboratory Director reviews laboratory data for quality and integrity, and she ensures that all personnel in the laboratory have received proper training and that they are following the procedures outlined in the Laboratory Manual and the Laboratory QA Manual.  She reviews training requirements and daily progress with the Laboratory Manager.  She orders supplies, materials, and equipment for the laboratory.  The Laboratory Director is available for on-call technical support at the gaging stations.

Sara Wilson and Jack Lewis, Data Analysts

The data analysts review and interpret data and compute suspended sediment load estimates.  All software used in the analysis was developed, and is maintained by, Redwood Sciences Laboratory, USDA Forest Service.  This software was designed specifically for data collected the Turbidity Threshold Sampling method.

Chris Mackelburg, Laboratory Manager

The Laboratory Manager will oversee the daily operation of the laboratory that includes sample processing, data entry, and training.  The Laboratory Manager will ensure that all equipment and instrumentation is functioning correctly and that the Laboratory Technicians have read and are following Laboratory and QA Manuals procedures and guidelines.  The Laboratory Manager will meet with the Laboratory Director on a regular basis to discuss processing goals and accomplishments, data quality, supply orders, staffing schedules, and training needs.  The Laboratory Manager proofreads data entered into databases against the original data sheets, clarifies ambiguous data with field personnel, and reviews field and lab data for QAPP objectives.

Points West, GPS Surveyor

The surveying position is held by a professional surveyor.  The surveyor will be responsible for supervisor all surveying and he will set all GPS control points.  The surveyor will use established surveying practices to ascertain all coordinates and elevations.

Laboratory Technicians

Laboratory Technicians process suspended sediment samples, measure laboratory turbidity, and enter laboratory data into the computer.

Randy Klein, Hydrologist

The hydrologist develops turbidity duration models, discharge rating curves, and analysis for comparing watersheds.

Rand Eads, Surface Water Specialist, RiverMetrics LLC

Drafts the original QAPP, submits draft QAPP for review, and prepares the final QAPP after addressing the reviewers’ comments.  Fabricates sampling equipment, provides training on field software and instrumentation, develops and writes field manuals, and installs equipment and instrumentation.

Station Masters, Jesse Noell and Clark Fenton

Supervise and participate in the maintenance and storm data collection at the six gaging stations.

Discharge Crew, Jesse Noell

The discharge crew is responsible for collecting discharge measurements at designated sites.

4.2 Quality Assurance Officer role

Tyler Ledwith, QA Officer

The QA Officer conducts lab and field certification, documents lab and field participant proficiency through checklists, and conducts periodic visits to observe lab and field techniques.  The QA Officer analyzes Quality Control checks, reviews approximately 10% of field and lab data for QAPP objectives.  The QA Officer also analyzes QC field and lab tests performed by the Project Director and Laboratory Director, respectively.  Results of these analyses and corrective actions are reported to the Project Director.

4.3 Persons Responsible for QAPP Update and Maintenance

Tyler Ledwith, QA Officer

Submits the final QAPP for signatures and distributes QAPP documents.  Changes and updates to this QAPP may be made after a review by the State Board’s Contract Manager and Quality Assurance Officer.  The QA Officer will be responsible for updating and maintaining the QAPP document throughout the life of the grant.

4.4 Organizational Chart

STATE WATER RESOUCES CONTROL BOARD, Bill Rey

QUALITY ASSURANCE OFFICER,

Tyler Ledwith

SALMON FOREVER

REDWOOD COMMUNITY ACTION AGENCY

PROJECT DIRECTOR, Jesse Noell

LAB DIRECTOR, Kristi Wrigley

STATION MASTERS

FIELD SURVEY

LAB MANAGER

# 5.0 Study Areas and Problem Statement

5.1 Problem Statement

Elk River and Freshwater Creek have been listed, or are in the listing process, as sediment-impaired watersheds under the Clean Water Act, Section 303d (CSWRCB, 1998).  Total Maximum Daily Loads (TMDL) has been developed for Elk River and they are currently under development for Freshwater Creek to identify and mitigate sediment impacts and to provide for attainment of Basin Plan water quality objectives.  In addition, Elk River and Freshwater Creek are identified as key watersheds for salmon production by the Northwest Forest Plan (USDA 1997).  Anadromous fish stocks on the North Coast of California have declined well below historical levels (Brown et al. 1994).  Increased sediment delivery to stream channels is thought to be a significant contributor to the reduction of fish populations.  Sediment can contribute to the decline of fish populations through several mechanisms including, but not limited to, clogging spawning gravel, impaired site feeding, reduced growth rates, and the reduction of habitat diversity through the loss of pool volumes and from aggradation in low gradient reaches (Chapman, 1988, Frissel, 1992, Newcombe and MacDonald, 1991).

5.2 Decisions or Outcomes

Efforts to restore the fisheries on Freshwater Creek and Elk River are planned and the effectiveness of these projects will benefit from monitoring turbidity, estimating suspended sediment loads (SSL), and monitoring changes to channel cross-sections.  Effective and accurate long-term monitoring using Turbidity Threshold Sampling (TTS) will allow assessment of restoration and remediation projects by measuring whether reduced suspended sediment loads and chronic turbidity levels have been achieved (Eads and Lewis, 2002, Klein 2003, Lewis and Eads, 2001, Lewis, 1996).

5.3 Water Quality or Regulatory Criteria

Three criteria form NORTH COAST REGION WATER QUALITY CONTROL BOARD 303(d) LIST UPDATE RECOMMENDATIONS, dated November 16, 2001, are relevant to this grant:

The water quality objective for sediment states:

“The suspended sediment load and suspended sediment discharge rate of surface waters shall not be altered in such a manner as to cause nuisance or adversely affect beneficial uses.”

The water quality objective for settleable material states:

“Water shall not contain substances in concentrations that result in deposition of material that causes nuisance or adversely affect beneficial uses.”

The water quality objective for turbidity states:

“Turbidity shall not be increased more than 20 percent above naturally occurring background levels.  Allowable zones of dilution within which higher percentages can be tolerated may be defined for specific discharges upon the issuance of discharge permits or waiver thereof.”

Criteria from the ADMINISTRATIVE PROCEDURES UPDATE APU 90-004, CLEAN WATER ACT  40 C.F.R. section 131.12 and STATE WATER BOARD RESOLUTION No. 68-16 antidegradation policy as are also relevant to this grant:

To implement the antidegradation policy the Regional Board should make findings when an existing discharge has reduced water quality, since an activity was last permitted and the reduction is not authorized by the permit.  If the Board finds that lowering of water quality is consistent with those conditions established in the State policy and the federal regulation, the findings should:

1.    Identify the pollutants that will lower water quality;

2.    Identify the socioeconomic and public benefits that result from lowered water quality; and

3.    Identify the beneficial uses that will be affected.

The baseline quality of the receiving water determines the level of water quality protection.  Baseline quality is defined as the best quality of the receiving water that has existed since 1968 when considering Resolution 68-16, or since 1975 under the federal policy, unless subsequent lowering was due to regulatory action consistent with State and federal antidegradation policies.

# 6.0 Project Description and Schedule

6.1 Work Statement and Produced Products

Salmon Forever operates continuous TTS stations (Lewis, 1996) to develop relationships between turbidity and suspended sediment concentration and to estimate both storm and annual suspended sediment loads.  The instrumentation at TTS stations includes an in situ turbidity probe, submersible pressure transducer for stage measurement, and an ISCO automatic pumping sampler with 24 sample bottles.  Sample collection is controlled by an algorithm that triggers sampling based on pre-established rising and falling turbidity thresholds.  Discharge rating curves are created for each TTS Station following the US Geologic Survey guidelines (Nolan and Shields, 2000, Buchanan and Somers, 1969, Harrellson, 1994).  Salmon Forever will employ the TTS method to collect turbidity and suspended sediment data in the Freshwater Creek and Elk River watersheds for Hydrologic Years (HY) 2008 and 2009.  All work is planned to start in the fall of HY 2008.

Several water surface elevation stations will placed on the lower Elk River to develop a data set for flood modeling.  The stations will consist of a data logger, pressure transducer, and battery.  Data will be collected every 10 minutes.

A Survey Crew will establish channel cross-section benchmarks throughout the two watersheds using a total station.  Benchmark elevations will be referenced to control points.  All channel cross-sections will be periodically measured throughout the two-year study.

A Discharge Crew will measure flows and update discharge rating curves at all established sites.

Soil pins will be used to document streambank and channel scour and fill.  These are measured annually.

Depth-integrated samples (DIS) will be simultaneously collected with TTS samples at gaging stations to determine the cross-sectional variability in sediment transport.  Depth-integrated sampling provides a mean flow-weighted sample by use of an isokenetic sampler such as a DH-48 or a D-74 (Edwards and Glysson, 1999).  Suspended sediment samples collected, via a pumping sampler, are considered point samples.  Point samples may not be representative of the cross-sectional average concentration because adequate mixing depends on grain size, local velocities, pumping sampler head height, and distance from the nearby tributaries.  Regression analysis is used to detect point sample bias.  Generally, point data are not corrected unless there a significant error, usually greater than 10%.

The Sediment Laboratory, operated by Salmon Forever, processes samples from the two watersheds to determine SSC.

6.2 Constituents to be Monitored

Water discharge measurements will be periodically collected in the field with a Price AA current meter or during low flows with Price Pigmy meter.  In situ turbidity will be measured and recorded every 10 minutes at four gauge sites with a DTS-12.  Suspended sediment concentration will be determined in the laboratory.  Turbidity, with three replicate sub samples, will be measured in the laboratory with a Hach 2100P.  Laboratory turbidity will not be used to compute sediment loads.

6.3 Project Schedule

Table 6-1.  Project schedule and deliverables.

| Activity | Anticipated Start Date | Completion Date | Deliverables | Deliverable Due Dates |
| --- | --- | --- | --- | --- |
| Start Project | 7/1/07 | 3/1/2010 | NA | NA |
| Equipment Installation | 9/15/07 as funding permits | 4/30/2008 | GPS coordinates | 12/31/07 |
| Training | 9/15/07 | 3/1/2010 | NA | NA |
| Sample Collection | 10/1/07 | 3/15/2010 | NA |  |
| Discharge Measurements | 11/1/07 | 3/15/2010 | Results Table, Rating Curve | 5/31/08; 5/31/09 |
| Surveying | 7/1/07 | 3/15/2010 | Data, Charts | 10/31/08; 10/31/09 |
| Lab Analysis | 12/1/07 | 3/15/2010 | Data Summary/ Analysis | 10/31/08; 10/31/09 |
| Data Analysis | 5/1/08 | 3/15/2010 | Data Summary/ Analysis | 10/31/08; 10/31/09 |
| Summarize Data |  | 4/1/2010 | Report of Data Summary & Analysis | 11/31/08; 11/30/09 |
| Annual Reports |  | September 1 2008, 2009 | Report | September 30 2008, 2009 |
| Final Report |  | March 1, 2010 | Report | March 7, 2010 |

6.4 Geographic Setting

The six gaging stations, two in Freshwater Creek, and four in Elk River flow into Humboldt Bay, and are listed in table 6-2 below:

Table 6-2.  Gaging site locations.

| Site | Latitude | Longitude | Area (mi2) | Date Installed |
| --- | --- | --- | --- | --- |
| FTR | 40°45.205’ N | 124°03.017’ W | 13.12 | 1999 |
| HHB | 40°45.878’ N | 124°03.953’ W | 27.78 | 2004 |
| KRW | 40°41.909’ N | 124°08.513’ W | 19.30 | 2003 |
| SFM | 40°42.077’ N | 124°08.936’ W | 22.19 | 2003 |
| MEC | 40°42.37’ N | 124°09.47’ W | 45+_ | 2008 |
| MEB | 40°43.28’ N | 124°10.47’ W | 50+_ | 2008 |

6.5 Constraints

Flood events can block roads and make access to some sites difficult.  Large storm events can cause equipment to malfunction when organic debris and sediment affect measurement devices and result in the loss of data.  Underwater electronic sensors suffer from deterioration leading ultimately to sensor failure.  Aggradation of the channels or streambank failure may require repositioning the sampling equipment so that it will function correctly.

# 7.0 Quality Objectives and Criteria for Measurement Data

Strict adherence to Standard Operating Procedures (SOPs are listed in the Appendix) ensures data quality by explicitly defining collection and analysis methods.  The precision, accuracy, and measurement range for the instrumentation are presented below.  Completeness for all samples collected will be greater than 90%.

Table 7-1.  Precision, accuracy, and measurement range for sensors and instruments

| Parameter | Device | Measurement Method | Precision | Accuracy or Non-linearity | Recovery | Measurement Range | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Turbidity | Hach 2100P | Laboratory | + 10% or 0.1 which ever is greater | + 10% or 0.1 which ever is greater § | NA | 0-40 NTU direct measurement; 40-999 NTU by dilution with turbidity free water | 90% |
| Turbidity | DTS-12 | In situ | NA |  2.0%  4.0%§ | NA | 0-500 NTU 500-1600 NTU | 90% |
| SSC Laboratory | - µ Glass filter | Gravimetric |  5.0 % |  2.0%§ | ±5% | 0.00001-2.0 g/L | 90% |
| Velocity | Floating Object | Float |  8.0% | 1.0 ft/sec. | NA | 0-10 ft/sec | 90% |
| Velocity | Price AA | Meter |  8.0% |  8.0%§ | NA | 0.25 - 8.0 ft/sec | 90% |
| Depth | Staff Plate | Visual |  5.0% | 0.02 ft | NA | NA | 90% |
| Depth | Druck 1830 | In situ | See accuracy | 0.1% FS BSL | NA | Device specific | 90% |

§ The accuracy or non-linearity for these parameters is a function of the range.

The accuracy of turbidity measurements and sediment load estimates is reliant on a number of factors:

- Selection of the appropriate turbidity thresholds in the TTS data logger program

- Variability in the turbidity record caused by air bubbles, fouling, and sensor position

- Proximity of the pumping sampler intake to the turbidity sensor

- Pumping sampler intake mounting to eliminate dips and cross-contamination

- Programming the pumping sampler to pre-rinse before sampling

- Sediment particle sizes in transport, pumping sampler lift, and stream velocity

- Organic transport and sediment color

- Sediment distribution in the sampling cross-section

- Sample holding time in the laboratory (minor effect)

- Choice of analysis models and methods

The combined error of all these parameters would be difficult to predict and the error will vary with the conditions.  Output statistics are available from the analysis models.  Although there are guidelines, the data analyst must use professional judgment when selecting models, reconstruction methods, and when defining storm events.  With the advent of efficient and accurate methods to estimate SSL, such as TTS, the largest source of error is often in the discharge rating.  Poor quality rating curves can create large uncertainties in the SSL estimates because the load is a product of discharge and SSC.  Careful assessment of the existing rating curves will allow the Discharge Crew to populate missing rating data and to make repeat measurements to detect rating shifts, and to verify relationships.

# 8.0 Special Training Needs and Certification

8.1 Specialized Training or certification

All participants that collect or analyze water quality samples receive training and certification for the specific tasks they will perform by the QA Officer in conjunction with the Laboratory Director or Project Director.  The goal of training and certification is to provide a safe working environment and to meet the Data Quality Objectives (DQO).  Field training takes place at established sampling sites and laboratory training is conducted at the Sediment Laboratory.

Training topics include:

- Safety

- Data recording and field forms

- Water quality sampling

- Discharge measurements

- Equipment maintenance & operation

- Laboratory sample processing

- Cross-section and scour pin surveys

- Data entry

8.2 Proficiency Checks

Proficiency checklists are used for evaluating the performance of participants.  The Project Director, Laboratory Director, and QA Officer use these checklists during training to document participant’s proficiency at identified tasks.

8.3 Safety

Safety procedures for sampling during storms or hazardous conditions are explained at every training session.  Collecting data during storm events and driving in inclement weather are the primary hazards that participants encounter in the field.  Risk from mechanical injuries can occur while operating the sampling crane.  Lifting injuries can arise from positioning sounding weights and transporting sample bottles.

Training encourages commitment and responsibility of participants in addition to imparting awareness of potential errors.  Requirements for participants include good physical health, the ability carry out and understand sampling and laboratory procedures, and availability to perform field sampling data analysis during non-standard work schedules.

# 9.0 Documents and Records

#### 9.1 QA Project Plan Distribution

The QA Officer will ensure that the most current QAPP version is available to all sponsoring and cooperating organizations involved in this study.  The individuals on the Distribution List will receive copies of the QAPP after each revision.  Current versions of the QAPP will be available upon request to any individual or organization.

#### 9.2 Field Documentation and Records

Field data are initially recorded on the appropriate data sheets or collected electronically by data loggers.  All field data sheets and notebooks are made by Rite-in-the-Rain and they are water resistant.  Copies of data sheets and notebooks are made and stored in the Sediment Lab.  Back up data logger files are kept on a minimum of two separate media (e.g. hard drive and CD).

#### 9.3 Laboratory Documentation and Records

Laboratory data documentation and recording varies with the analyses being performed.  Sample sign-in sheets, turbidity determination, and suspended sediment concentration data is recorded in pen on all appropriate data sheets (see the Appendix).  Suspended sediment calculations are performed and maintained in an Excel spreadsheet and are written on paper worksheets.

Original datasets are kept on the Laboratory Director’s computer.

All data summaries and databases developed from raw parameter measurements are maintained in MS Word and Excel formats.  Paper copies are maintained on file.  All data sheets include the hydrologic year, initials of the person entering data, date of entry and, date of copying.  Sign-in sheets and filter tare sheets are numbered sequentially throughout the hydrologic year.  Laboratory data sheets are filed chronologically and given sequential numbers at the end of each hydrologic year.

9.4 Retention

Originals of all field and laboratory data sheets, QA/QC analyses, and computer databases will be kept in the Sediment Laboratory.  Copies of field and laboratory data sheets, QA/QC analyses, and computer back-up disks will be kept in the Salmon Forever office.  Salmon Forever will maintain hard copies of all data sets and computer back-up disks for at least 23 years.  Original ISCO Automatic Sampler field sheets will be stored for 23 years at the Sediment Laboratory.  Copies of these documents will also be given to Redwood Sciences Lab for safekeeping.

#### 9.4 Reports

A QA report will be prepared by August 1, of 2008 and 2009.  An annual project report will be completed by September 1, 2008.  A final report will include raw data, field data sheets, suspended sediment data sheets, equipment calibration sheets, laboratory data sheets, QA/QC results, turbidity records, discharge records, and suspended sediment load estimates.  The final report will be completed by March 1, 2010.

# Group B:  Data Generation and Acquisition

# 10.0 Sampling Process Design

The Monitoring Plan discusses sampling process design in detail.  Sampling sites are chosen to meet the following project goals:

- To develop predictive relationships between turbidity and suspended sediment concentrations,

- To develop turbidity duration data sets, and,

- To identify the effects of land use activities in the watershed by sampling above and below select locations within a watershed.

# 11.0 Sampling Methods

The SOPs (see the Appendix for a complete list of SOPs) contain detailed information on the methods and procedures used for sample collection and analysis.

# 12.0 Sample Handling and Custody

Samples are identified using unique stickers attached to bottles.  All sample bottles used for sampling are assigned a waterproof sticker with a unique Sample ID number.  The Laboratory Manager affixes labels to all empty sample bottles before they taken into the field.  In the field, samples are labeled according to the format in table 12-1:

Table 12-1.  Sample bottle label description.

| Label Code Description | Example |
| --- | --- |
| Two-digit Hydrologic Year | 07, 08, 09 |
| Two-character sample type: IS– ISCO sample DI – Depth-Integrated sample | 07IS1413 07DI0049 |
| Four-digit unique sequential number for each sample type within the Hydrologic Year | 0001 – 9999 |
| TTS samples (additional label) HY (07,08,09) Station identifier (FTR, HHB, KRW, SFM) Bottle sequence (01-24) | 07KRW17 |

DI sample bottles are further labeled in the field with: observer initials, site, time, date, stage.  ISCO sample bottles are labeled before removal from the sampler.  After turbidity and SSC processing is completed the stickers are removed from the sample bottles.

No special handling is required for the samples.  After the samples are collected in the field, they are stored in cool location and are covered with black plastic to eliminate light.  Samples are processed in the order in which they arrive at the laboratory to reduce the holding times.  Laboratory turbidity is measured as soon as samples are logged-in at the laboratory.  Samples are then acidified to retard biologic growth by adding four drops of acid (1 part HCL to1 part water) to each sample.  Samples are held in the laboratory in a cool location and covered with black plastic.  Testing has been conducted by Redwood Sciences Laboratory, US Forest Service, to determine the effects of holding times on biologic growth and changes in sediment mass (see section 14).  Because no special disposal procedures are required samples are discarded after processing is complete.

No special chain-of-custody procedures are required for the samples.  The date and time that the TTS samples were collected are recorded in the electronic file.  The DI sample collection information, and TTS samples that are removed from the ISCO sampler, are recorded in the field log.  In addition, the date, time, and bottle numbers of the TTS samples that were removed from the ISCO sampler are recorded on the TTS Field Form (see the Appendix).  The TTS field form is used to record current hydrologic conditions.  The field log describes the disposition of the samples, that is, whether they are in transit or being held at a particular location.  The log includes the date and time the samples delivered to the laboratory.  Once the samples arrive at the laboratory their identifying information is recorded on the “Sample Sign-in Log” (see the Appendix).  Laboratory processing forms track the samples throughout the analysis process (see the Appendix).

The Project Director is responsible for custody of samples in the field and the transport of the samples to laboratory.  Station Masters may assume the field responsibilities at various times.  The Laboratory Director and Laboratory Manager are jointly responsible for the custody of samples once they arrive at the laboratory.

# 13.0 Analytical Methods

Analysis is not performed on samples in the field.

Analytical laboratory procedures use these references to determine suspended sediment concentration and turbidity, table 13-1:

Table 13-1.  Laboratory analytical methods.

| Analysis | Matrix | Reporting Units | Method | Target Reporting Limit |
| --- | --- | --- | --- | --- |
| Suspended Sediment Concentration | Water | mg/L | RSL 2007 ASTM 2000 D3977 Gray et al 2000 | 0.5 mm |
| Turbidity | Water | NTU | EPA 180.1 | 0.5 NTU |

The Laboratory Director, in discussion with the QA Officer, is responsible for correcting any failures in the analytical process.  Detailed information on the corrective actions and any samples affected shall be kept in the laboratory records.

# 14.0 Quality Control

#### 14.1 Field Sampling

Depth-integrated manual sampling is performed concurrently with user-initiated pumped samples.  This provides a method to correct for possible bias in TTS point samples.  Depending on the channel geometry, mixing of sediment from nearby tributaries, and location of the pumping sampler intake, the point samples collected by the pumping sampler, under TTS control, may not be representative of the cross-sectional flow-weighted suspended sediment concentration.  The number of sample pairs required to determine if a bias correction is required is site dependant, but generally 15 to 20 samples distributed over range of measured turbidities, and during span of the project, is adequate.

Because of the naturally high variability in suspended sediment transport, and the limited number of available sampling bottles in the pumping sampler, it is not feasible to collect duplicate samples.  Turbidity measurements are made every 10 minutes.  One hundred turbidity measurements are processed in a five-second period and the median value is recorded in the data logger file (the median is a more robust statistic than the mean when some individual measurements may be extreme values).

#### 14.2 Laboratory Analysis

The Quality Assurance Officer is responsible for implementing, recording, and analyzing the quality control measures undertaken to ensure data quality objectives are met.  Quality Control measures for each sampling procedure are detailed in the SOP (SOP 4, Suspended Sediment Laboratory Procedures).  In general, quality control will represent 10% of the data collected in this study.  Results of quality control analyses and corrective actions are reported to the Project Director and described in the annual report.

Table 14-1.  Laboratory analysis.

| ### Group | ### Parameter | ### Quality Control | ### Instrument Calibration Frequency |
| --- | --- | --- | --- |
| Laboratory Measurement | Turbidity | 3 replicates from each sample | Follow manufacturer’s guidelines |
| Laboratory Analysis | Suspended Sediment Concentration | Filter blanks 1 in 20 Sample splits to RSL Spiked samples | Annual calibration of analytical balance with traceable standards Check weight 1 in 10 |

14.3 Estimates of Field and Laboratory Sampling and Processing Errors

The TTS sampling method is fully automated and only requires field service to exchange sample bottles and batteries, and to download the data to a laptop.  Field personnel are at the gage sites for brief periods during storm events and during routine site maintenance visit.  Most of the standard procedures and holding times that commonly employed for manual sampling are not applicable with the TTS method.

Both Freshwater Creek and Elk River are impaired for sediment, and consequently SSC values are normally high and concentrations may become extreme during large events.  Because of the high concentrations and sampling errors (see below), errors associated from bacterial or algae propagation in the sample bottles is considered insignificant under all conditions except the lowest concentrations.  Holding times for samples collected in the field may range from 1 to 7 days before the samples are transported to the laboratory.  Because turbidity is measured in real-time in the field (under data logger control) laboratory turbidities are not used for developing a relation with SSC.  Errors in SCC from the dried mass of bacteria or algae contamination are thought to be insignificant when concentrations range from the hundreds to the thousands (mg/L), when taking into consideration all possible errors.

A study at Redwood Sciences Laboratory, USDA Forest Service, gives some incite to the possible errors associated with holding times in the laboratory (personal communication, Jack Lewis, Redwood Sciences Laboratory):

Omitting questionable data, only 2 of 4 analyses showed a significant effect.  The minimum increase in measured SSC was 7 mg/L.  With all the data, the average increase was 18 mg/L.  Here are the SSC means from week 0 and week 12:

Week 0      Week 12      Change 
storm 1 rising               95               109               14%
storm 1 falling             203              224               21%
storm 2 rising              372              397               25%
storm 2 falling             119              132               13%

The errors listed below are general and they are the lowest errors to be expected for “good” measurements, but site-specific errors may vary considerably based on the stream and techniques used (personal communication, John Gray, Sediment Specialist, US Geologic Survey).

- Water discharge, +/- 5%, errors can be much larger if the channel control is not stable, or when field procedures are not strictly followed

- Depth-integrated samples, +/- 5% for medium sands and finer

- Laboratory processing of SSC (data from the USGS National Laboratory QA Program; Gray and Glysson, “USGS Sediment Data Collection Techniques" a training course in held Castle Rock, WA, 2008):

| SSC | Material Size | Variance |
| --- | --- | --- |
| 10 | Fines | ± 50% |
| 1,000 | Fines | ± 5% |
| 1,000 | Coarse | ± 40% |

The magnitude of load computation errors depends on the analysis software used, the analyst’s experience, and the quality of the data.  It is likely that the errors are in the range of +/- 10 to 25% [R. Eads].

# 15.0 Instrument Testing, Inspection and Maintenance Requirements

All equipment is inspected and maintained to EPA and manufacturer specifications.  Records of maintenance and calibration are kept for all appropriate equipment.  The Laboratory Manager maintains these records to schedule maintenance on all equipment.  All records and laboratory equipment will be kept at the Sediment Lab, in additions to spare parts.  If equipment does not meet specifications, or it is not working properly, it shall not be used until inspected by the QA Officer and acceptance, repair, or replacement has been documented.  Table 15-1 summarizes the inspection frequency and performance assessments used to identify equipment malfunctions.

Table 15-1.  Equipment and instrument inspection and maintenance.

| Equipment | Inspection Frequency | Type of Inspection or Assessment | Inspector |
| --- | --- | --- | --- |
| Balance | Each use | Weigh check weights every 10th weigh | Laboratory Director or Lab Technician |
| HACH Turbidimeter | Each use | Check for proper operation | Lab or Project Director |
| ISCO sampler | Each site visit | Check for proper operation, does display bottle number match data logger number | Station Master |
| Data logger | Each site visit | Check for correct sensor data outputs | Station Master |
| Druck 1830 Pressure transducer | Each site visit | Check for correct data output and compare to staff plate | Station Master |
| DTS-12 and OBS-3 Turbidity sensors | Each site visit | Check for proper data output | Station Master |
| TE 525I Tipping Bucket Rain gauge | Each site visit | Check for debris and proper operation | Station Master |
| DH-48 and DH-49 samplers | Each use | Visual inspection | Project Director or Field technician |
| Price AA and Pygmy Current Meters | Each use | Visual inspection and spin test | Project Director or Field Technician |

# 16.0 Instrument/Equipment Calibration and Frequency

Calibration frequency and methods are shown below.  All equipment calibration records are maintained by the Laboratory Manager.  Each piece of equipment has an identifying number that is linked to calibration records.  Table 16-1 summarizes the instrument calibration schedule.

Table 16-1.  Instrument calibration schedule.

| Instrument | Calibration Frequency | Type of Calibration | Responsible Party |
| --- | --- | --- | --- |
| Balances (analytical and top-loading) | Annual | Quality Control Services Standard Wts No. 349-B traceable to NIST Standard Weights | Quality Control Services 2340 SE 11th Ave Portland OR 97214 |
| HACH laboratory turbidity meters, SOP 2 | Quarterly | Stablcal calibration solution | Laboratory Manager |
| In situ pressure transducers, SOP 7 | Beginning of season | RSL protocol | Lab or Project Director |
| In situ turbidity sensors, SOP 6 | Beginning of season | RSL protocol or factory calibration | Lab or Project Director |
| Tipping Bucket Rain gauge, SOP 8 | Beginning of season | RSL protocol | Lab or Project Director |

# 17.0 Inspection/Acceptance of supplies and Consumables

Supplies will be examined for damage as they are received.  The following supplies will receive additional inspections as follows:

- Each 50 mm, 1-µ glass filter, is examined for holes when it is placed on the vacuum manifold.

# 18.0 Non-Direct Measurements

The only non-direct measurements are from the Salomon Forever’s database of gaging station data (streamflow, turbidity, SSC, rainfall, temperature, discharge ratings, cross-sections, and suspended sediment loads) from prior years (same sites as this project).  The database is maintained in accordance with Salmon Forever’s policy as stated earlier.  The data will be reviewed against the data quality objectives stated in section 7 and only that data meeting all of the criteria will be used in this project.

Participants associated with Salmon Forever collect most of the data required to analyze the turbidity response and sediment transport.  At a few sites, Humboldt State University (HSU) students and professors are also conducting monitoring projects, primarily continuous stage measurements and rating curve generation.  Salmon Forever requests copies of data and analysis from cooperators as the need arises.

Additionally, Salmon Forever also maintains a collection of maps and aerial photographs of sampling sites and target watersheds.  These materials are acquired from the USGS and other appropriate agencies.

# 19.0 Data Management

Data will be maintained as described in Section 9 above.  Sample information and data are all recorded on standardized field and laboratory data forms (see the Appendix).  The Project Director is responsible for checking and copying field data sheets.  The Laboratory Manager is responsible for checking and copying lab data sheets and delivering them to the Project Director.  Original laboratory data sheets are kept in the Sediment Laboratory.  Reports and data are transferred to MS Excel spreadsheets and MS Word documents, and copies kept at the Sediment Laboratory.

All electronic TTS data are examined, rated, and assigned field and laboratory codes pertaining to the quality of data in Table 19-1 below:

Table 19-1.  TTS stage and turbidity quality codes.

Data handling forms and software includes data sheets, MS Excel spreadsheets, analysis programs written in R and developed by RSL, hand calculators, and Microsoft Word documents (see the Appendix).

# Group C: Assessment and Oversight

# 20.0 Assessment and Response Actions

Project officers meet every 2 months to compare progress, to discuss and resolve problems that they may have encountered, and to address any issues brought to their attention by the results of internal QA checks.  These meetings identify problems with sampling procedures or logistics in the field.  Project officers will discuss difficulties encountered in specific situations, adopt corrective actions, and implement appropriate modifications for standardizing methods between participants.

The Project Director, QA Officer shall observe each participant at the beginning of the project, and thereafter at least twice a year, conducting sampling.  Proficiency checklists (see the Appendix) will be used to provide a written record and evaluation of participant performance.  All participants are required to pass proficiency evaluations during training.  If participants do not meet the proficiency criteria, they will receive additional training until they are proficient or they will not be utilized in this study.  Participants are required to perform all sampling procedures correctly, before their data is included in any databases or is used for analyses.

Laboratory QA procedures are detailed in the SOP (SOP 4, Suspended Sediment Laboratory Procedures).  Either the QA Officer or Laboratory Manager conducts QA laboratory procedures.  The Laboratory Manager trains lab technicians before they begin conducting sample processing and observes their proficiency on the job until they are certified to work independently.  Technicians work under direct supervision for a minimum of 2 sessions.  Certification using the proficiency checklists for turbidity and SSC determination and filter weighing are conducted for all lab technicians at the beginning and mid-point of each sampling season.  The Laboratory Director reviews technician data for errors and incomplete data entry.  The Laboratory Director is responsible for implementing these assessments, correcting technician deficiencies, and keeping the checklists on file in the lab.  Results of assessments and certifications are reported to the Project Director.

Technical system audits provide an external review of the research and QA activities.  External personnel from EPA, NCRWQCB, RSL, or HSU will audit this project during the field season.  Findings will be discussed with the participants and summarized in audit reports submitted to EPA, NCRWQCB, RSL, and HSU.

The objectives of field or laboratory visits are to:

- Observe implementation of field methods by the field crews;

- Assess personnel performance, and examine equipment and procedures;

- Evaluate training methods;

- Assess consistency of participants in implementing field methods;

- Answer questions arising about sampling design or methods; and

- Determine whether DQOs are being met from review of quality assurance data

If deficiencies or problems are identified that are determined to be jeopardizing data quality, outside assessors will make recommendations to the Project and Laboratory Directors.  The external reviewers have the authority to stop the activity until the problems are resolved.

# 21.0 Reports to Management

An annual report will be produced and distributed in August of each year.  The Project Director is responsible for report production and distribution.  Reports will be forwarded to the county, state, and regional agencies, and other members of the Advisory Panel.  Reports will contain data analysis for the previous year’s sampling, an update on project status and findings to date, participant highlights, results of quality assessment audits and internal assessments, and it will identify any significant QA problems and their solutions.  The Project Director will incorporate the recommendations in this report by implementing required changes to the study for the next sampling season.

Progress reports will be produced quarterly.  The first report will be submitted four months after the grant is authorized.

In addition to the annual project report, an annual QA report is also be prepared by the QA Officer.  The QA Report summarizes the outcome of all quality assurance efforts undertaken for the sampling season and makes recommendations for improving activities for the next year.  The QA Report specifically addresses data quality and information management by:

- Evaluating all QA data and sampling;

- Summarizing data entry errors and describing any difficulties with data;

- Evaluating data entry completeness;

- Documenting data management activities, including the content and location of project notebooks (field, laboratory, data management) and data sheets.

The QA Report also summarizes the results of quality assurance activities identifying the greatest sources of error and evaluating the SOPs, DQOs, and training effectiveness.

A final report is due 1 month following the completion of the grant.

# Group D: Data Validation and Usability

# 22.0 Data Review, Validation and Verification Requirements

Data generated by project activities will be reviewed against the data quality objectives cited in Element 7 and the quality assurance/quality control practices cited in Elements 14, 15, 16, and 17.  Data will be separated into three categories: data meeting all data quality objectives, data meeting failing precision or recovery criteria, and data failing to meet accuracy criteria.  Data meeting all data quality objectives, but with failures of quality assurance/quality control practices will be set aside until the impact of the failure on data quality is determined.  Once determined, the data will be moved into either the first category or the last category.

Data falling in the first category is considered usable by the project.  Data falling in the last category is considered not usable.  Data falling in the second category will have all aspects assessed.  If sufficient evidence is found supporting data quality for use in this project, the data will be moved to the first category, but will be flagged accordingly.  The exception to these guidelines concerns the holding times for suspended sediment samples (see previous discussions).

# 23.0 Validation and Verification Methods

Once lab data has been entered into the database, the Laboratory Manager verifies 100% of the data.  Errors in data entry are corrected.  Outliers and inconsistencies are flagged for further review or discarded.  Problems with data quality will be discussed in the in the quarterly and annual reports.  The Project Manager will randomly verify 10% of the lab data for accuracy and completeness and report the results to the QA Officer.

Following data entry and review, the QA Officer evaluates quality control data using appropriate techniques such as graphical comparisons and statistical analysis.  Results of the QA Officer findings are reported in the annual QA report.

# 24.0 Reconciliation with Data Quality Objectives

Calculations of precision, accuracy, and completeness will be made and included in the annual report.  If data quality indicators show that sampling methods are not meeting the project's specifications, the cause of failure will be evaluated and corrective action implemented.  If the cause is found to be equipment failure, calibration, or maintenance, techniques will be reassessed and modified.  If the problem is determined to be a sampling error, participants will be retrained.  Any limitations on data use will be detailed in the final report and other documentation as needed.

If failure to meet project specifications is found to be unrelated to equipment, methods, or sample error, specifications may be revised for the next sampling season. Revisions will be submitted to the RSL, EPA and NCRWQCB quality assurance officers for approval. Limitations on the use of data will be reported in the annual QA report.

# References

American Society for Testing and Materials (ASTM), 2000, Standard test methods for determining sediment concentration in water samples: D3977-97, vol. 11.02, Water (II), 395-400.

Brown, L. R., P. B. Moyle, and R. M. Yoshiyama. 1994. Historical decline and current status of coho salmon in California. N. Am. J. Fish. Manage. 14:237-261.

### Buchanan, T. J.; Somers, W.P. 1969. Discharge measurements at gaging stations. U.S. Geological Survey, Techniques of Water-Resources Investigations, Book 3, Chapter A8, vii, 65 p. :ill. ;26 cm. Reprinted in 1984.

California State Water Resources Control Board. http://www.sqrcb.ca.gov/tmdl/303d_lists.html.  Web page last updated 1/16/01.

Chapman. D.W. 1988. Critical Review of Variables Used to Define Effects of Fines in Redds of Large Salmonids. Transactions of the American Fisheries Society. 117: 1-21.

Edwards, T.K.; Glysson, D.G. 1999. Field methods for measurement of fluvial sediment. Techniques of Water-Resource Investigation Report Number 03-C2. US Geologic Survey.

Eads, Rand, and Jack Lewis. 2002. Continuous turbidity monitoring in streams of northwestern California. In: Turbidity and other sediment surrogates workshop (ed. by G.D. Glysson & J.R. Gray). 30 April - 02 May 2002, Reno, Nevada. 3 p.

Frissell, C.A. 1992. Cumulative effects of land use on salmonid habitat on southwest Oregon streams. Ph.D. thesis, Oregon State University, Corvalis, OR.

Gray, JR, Glysson, GD, Turcios, LM, and Schwarz, GE., 2000, Comparability of suspended-sediment concentration and total suspended solids data: USGS Water-Resources Investigations Report 00-4191.

Harrellson, C. C.  1994. Stream Channel Reference Sites: An Illustrated Guide to Field Technique: USFS, Rocky Mountain Forest and Range Experiment Station, General Technical Report RM-245.

Klein, Randy. 2003. Duration of turbidity and suspended sediment transport in salmonid-bearing streams, North Coastal California. A report to the USEPA Region 1X.

Lewis, Jack, and Rand Eads. 2001. Turbidity threshold sampling for suspended sediment load estimation. In: Proceedings, 7th Federal Interagency Sedimentation Conference, 25-29 Mar 2001, Reno, Nevada. [1824 KB]

Lewis, Jack. 1996. Turbidity-controlled suspended sediment sampling for runoff-event load estimation. Water Resources Research 32(7): 2299-2310.

Lewis, Jack. 1996. Turbidity-controlled suspended sediment sampling for runoff-event load estimation. Water Resources Research 32(7): 2299-2310.

Newcombe, C.P. and D.D. MacDonald. 1991. Effects of Suspended Sediments on Aquatic Ecosystems. North American Journal of Fisheries Management. 11: 72-82

Nolan, M and R. Shields. 2000, Measurement of stream discharge by wading.  WRIS 00-4036, version 1.1. USDI, USGS

North Coast Regional Water Quality Control Board.  1993.  North Coast Basin Plan. December 1993. http://www.swrcb.ca.gov/rwqcb1/download/basin_plan.pdf

Redwood Sciences Laboratory. 2007.  Sediment Laboratory Procedures. USDA Forest Service, Arcata, CA. (http://www.fs.fed.us/psw/topics/water/tts/index.shtml#tech).

USDA Forest Service. 1997.  1997 Northwest Forest Plan – An Ecosystem Management Approach.  .

USEPA.  1996.  The Participant Monitor’s Guide To QUALITY ASSURANCE PROJECT PLANS

Office of Wetlands, Oceans and Watersheds  EPA 841-B-96-003.  September 1996. http://www.epa.gov/OWOW/monitoring/participant/qapp/vol_qapp.pdf

US Geologic Survey, Department of Interior, Techniques of Water-Resources Investigations of the USGS:

Stage Measurements at Gaging Stations Book 3 Chapter A7

Discharge Measurements at Gaging Stations Book 3 Chapter A8

Laboratory Theory and Methods for Sediment Analysis Chapter C1 Book 5

Field Methods for Measurement of Fluvial Sediment Chapter C2 Book 3

Surface Water Techniques:  Discharge Ratings at Gaging Stations – Hydraulic Measurement and Computation Book 1 Chapter 12 1965

# Acknowledgements

Salmon Forever thanks the North Coast Regional Water Quality Board, and State Water Resources Control Board for providing funding for this project under the 2005 -2006 Consolidated Grants, Proposition 50 Coastal Nonpoint Source Pollution Control Program, Agreement No 06-289-551-0.  We also thank Redwood Community Action Agency (RCAA) for administering this project (Contract 1SF07_952, Humboldt Bay Sediment Reduction and Salmon Habitat Improvement Project, HB SedSHIP.

Salmon Forever acknowledges support from the Technical Advisory Committee, the local community and landowners, local and regional organizations, businesses, and scientists from state and federal agencies.

# Appendix

A-1.  Standard Operating Procedures (SOPs)

| SOP Number | SOP Title |
| --- | --- |
| 01 | Turbidity Threshold Sampling Field Methods |
| 02 | Laboratory Turbidity Measurements |
| 03 | Water Discharge Measurements |
| 04 | Suspended Sediment Laboratory Procedures |
| 05 | ISCO Automatic Pumping Sampler |
| 06 | In Situ Turbidity Sensor |
| 07 | Electronic Stage Measurement |
| 08 | Tipping Bucket Rain Gage |
| 09 | Depth-Integrated Sampling |
| 10 | Scour Pins |
| 11 | Cross-Section and Longitudinal Surveys |
| 12 | Electronic Water Surface Measurements |

A-2.  Forms, tables, and worksheets used in the field and laboratory

A-2-1		Depth table for discharge measurements

A-2-2		Hach 2100N Calibration sheet

A-2-3		Hach 2100P Calibration sheet

A-2-4		Price AA current meter calibration sheet

A-2-5		Price Pigmy current meter calibration sheet

A-2-6		TTS field form

A-2-7		Data dump log sheet

A-2-8		Balance check weight log

A-2-9		Druck pressure transducer calibration sheet

A-2-10		Initial filter weight sheet

A-2-11		Instrument maintenance sheet

A-2-12		Timed pump sample log

A-2-13		TTS station parameter settings

A-2-14		Laboratory sample sign-in sheet

A-2-15		Suspended sediment calculation worksheet

A-2-16		Turbidity dilution worksheet

A-2-17		Laboratory SSC processing sheet