STATE OF CALIFORNIA

DEPARTMENT OF FISH AND GAME

OFFICE OF SPILL PREVENTION AND RESPONSE

LABORATORY QUALITY ASSURANCE

PROGRAM PLAN

California Department of Fish and Game

Fish and Wildlife Water Pollution Control Laboratory

2005 Nimbus Road

Rancho Cordova, CA 95670

Section No.     1

Revision No.     6

Date     16 May 03

Page      1 of 1

SAMPLING AND ANALYTICAL ACTIVITIES

State of California

Department of Fish and Game

Office of Spill Prevention and Response

Scientific Program

Fish and Wildlife Water Pollution Control Laboratory

Approvals:

OSPR Scientific Program Chief:

Ken Mayer

OSPR Scientific Program Asst. Chief	__________________________

and Laboratories Manager:			John Turner

Laboratory Director,				__________________________

Fish and Wildlife Water Pollution 		David Crane

Control Laboratory

OSPR Quality Assurance			__________________________

Officer (Acting):  				Tom Lew

Contracting Agency				__________________________

Project QA Officer:

DEPARTMENTAL QUALITY ASSURANCE

PROGRAM POLICY

The Fish and Game Departmental (DFG) quality assurance program describes the requirements, controls and responsibilities for implementation of quality assurance principles specified in applicable regulations, codes, and standards applied to the environmental laboratory activities.  The program begins with quality assurance training for all new employees, and an orientation to the Departmental quality assurance/quality control practices.  The importance of quality assurance is recognized by Department management and is documented within the Office of Spill Prevention and Response for all laboratory operations.

The primary commitment of the Departmental Quality Assurance/Quality Control Program is to implement the program activities and requirements committing time and resources ensuring that data are as precise, accurate and complete as required by the data quality objectives of the projects involved.

Section No.     2

Revision No.     4

Date     03 Apr 00

Page      1 of 1

OFFICE OF SPILL PREVENTION AND RESPONSE

LABORATORY QUALITY ASSURANCE PROGRAM PLAN

TABLE OF CONTENTS

Section            Title

1		Title Page

2		Table of Contents

3		Quality Assurance Description

4		Quality Assurance Organization and Responsibilities

5		Data Quality Objectives and Assessment Methods

6		Sampling Procedures

7 		Sample Custody

8     		Calibration Procedures and Preventive Maintenance

9		Analytical Procedures

10		Data Reduction, Validation and Reporting

11		Internal Quality Control Checks and Frequency

12		Performance and System Audits

13		Preventive Maintenance

14		Routine Assessment of Data Precision, Accuracy, and Completeness

15			Corrective Action

16			Quality Assurance Communication with Management

17			Staff Training and Documentation

18 		Glossary

Appendices

A			Organizational Chart

B			Qualifications  and Specifications of Key Personnel

C			SOP’s

D			Sample Containers, Preservation and Holding Time

E			Forms

F			Analytical Methods and Reference Sources

G			Method Detection Limit Revision 1.11, Federal Register, Vol. 49, No. 209, Friday October 26, 1984.

Section No.     3

Revision No.     4

Date     03 Apr 00

Page      1 of 1

3.0    QUALITY ASSURANCE DESCRIPTION

3.1   Overview

The purpose of this document is to describe the State of California Department of  Fish and Game's Quality Assurance Program as implemented within the Office of Spill Prevention and Response (OSPR) Laboratories.  This program plan summarizes those quality assurance and quality control (QA/QC) elements which ensure the accurate and precise development of Department sampling and analytical results, as is consistent with project objectives.  The program plan has been designed to meet requirements of many projects and specifically addresses all elements of the Environmental Protection Agency Quality Assurance Management staff's "Interim Guidelines and Specifications for Preparing Quality Assurance Project Plans" QAMS-005/80.  This plan establishes the quality assurance and quality control procedures common to most of the Laboratory services. When necessary, particular project protocols or Standard Operating Procedures (SOPs) will be used to define any project-specific requirements.

3.1.1    Department Quality Assurance System

Quality assurance is a system for integrating the quality planning, quality assessment and quality improvement efforts of various sections to enable operations to meet specified project needs.  Quality assurance of field and laboratory systems is concerned with all activities that have an important effect on the quality of measurements as well as the establishment of methods and techniques to monitor the performance of these systems.  In addition, quality assurance is composed of those activities performed on a routine basis to gain an independent assessment of the operation and validity of the product.  In summary, quality assurance is an essential system of activities to provide the confidence that quality control methods are performing adequately.

3.1.2    Department Quality Control System

In contrast, quality control is the system of  activities which provide a quality product for a data user, consisting of internal laboratory operations which document product quality.

3.2   Summary

In summary, this Quality Assurance and Quality Control Program Plan is designed to satisfy the requirements and concerns of the analyst, management, and regulatory agencies concerned with the project.

Section No.     4

Revision No.     4

Date     03 Apr 00

Page      1 of 1

4.0    QUALITY ASSURANCE ORGANIZATION AND RESPONSIBILITY

4.1    OSPR Scientific Program Chief

The Scientific Unit Chief is responsible for administrative and financial oversight of all activities within the OSPR Scientific Unit.  Organizational chart can be found in Appendix A.

4.2    OSPR Scientific Program Assistant Chief

The Scientific Unit Assistant Chief is responsible for administrative and financial oversight of all activities within the OSPR Scientific Unit's laboratory system.

4.3    Laboratory Directors (Quality Control Coordinators)

Laboratory directors are designated for each of the laboratories within OSPR.  The laboratory directors are accountable for all operational activities, including examination of all analytical data, quality assurance parameters, and report preparation and review.

4.4    Quality Assurance Officer

The Quality Assurance Officer is responsible for the periodic introduction of performance evaluation samples where feasible, the evaluation of all sample logging/numbering procedures, and the evaluation of all quality control data.

4.5    Project/Section Leaders

Project Leaders are responsible for daily laboratory activities relating to their          individual project assignments.  Responsibilities include: making daily work assignments for laboratory staff, generation and review of data and preparation and initial review of all laboratory data reports.

  - Laboratory Staff

The responsibilities of the laboratory staff include sample container and glassware preparation, calibration standard and reagent preparation, sample preparation, analysis, and preparation of analytical reports with quality control data for the project/section leaders and laboratory supervisor.  Staff members will be familiar with all general laboratory procedures and quality assurance objectives.

  - Specifications for Key Personnel:  Appendix B

Section No.     5

Revision No.     2

Date     03 Apr 00

Page      1 of 1

5.0    DATA QUALITY OBJECTIVES AND ASSESSMENT METHODS

5.1    Overview

The primary data quality objective is to provide a product that fulfills all project and/or agency requirements.  The requirements for projects are established prior to their commencement.  In the absence of specific data requirements, standard methods or verified alternative protocols will be routinely applied.

5.2    Objectives

The data that is produced from the laboratory must be scientifically valid, defensible, comparable, and of known precision and accuracy.  Objective measures of data quality such as method blanks, duplicates, spikes, standards, and recoveries will be employed.  Acceptance limits will be established for data accuracy and precision.  Whenever possible, statistical methods such as confidence limits, significance tests and/or variability measures will be used to evaluate precision and accuracy of data as well as conformance to acceptance limits.  Control charts can be used to graphically represent measures of data quality.  Corrective action will be initiated when the quality of the data does not meet established quality standards.

  - Standard Operating Procedures

5.3.1    Standard Quality Control Procedures

Where appropriate, standard quality control procedures, data reduction, and reporting will be in compliance with requirements in Standard Methods for the Examination of Water and Wastewater, 18th edition (1992), with requirements in USEPA Handbook for Analytical Quality Control in Water and Waste Water Laboratories: EPA-600/4-79-019, and with the requirements in USEPA Test Methods for Evaluation of Solid Waste: Physical/Chemical Methods, SW-846, 3rd edition, 1986.

5.3.2    Standard Operating Procedures

Written standard operating procedures(SOPs) for receipt of samples, tracking of custody, sample preparation and analysis, use of equipment and instrumentation, and assembly of completed data shall be followed.  These SOPs shall include use of standard data logging formats, logbook/worksheet entry procedures, and other written or printed documents relevant to the samples.  These SOPs are available on request.

Section No.     6

Revision No.     3

Date     03 Apr 00

Page      1 of 3

6.0    SAMPLING PROCEDURES

6.1    Objectives

The reason for sampling and the parameters of concern for each sampling event establish the requirements for sample container type and preparation, sample amount and preservation, and the sampling technique.  Information on the sampling site is assembled so that a project work plan can be developed for the collection of representative samples.

6.2    Preparation for Sampling

Prior to conducting project field sampling operations, pre-cleaned sample containers and sampling devices are assembled along with the necessary equipment and portable instrumentation.  A team of trained personnel with appropriate protective equipment will then conduct actual sampling using established procedures.

6.2.1    Field Quality Measures

Field quality measures such as trip blanks for water control samples, field blanks,  duplicates, and background references are employed to assure data quality.  Sample filtration, when required, can be performed in the field.  Sample preservation is routinely provided by using sample containers with pre-added preservatives.  Sampling record sheets and chain-of-custody or record-of-custody forms are completed at the time of sampling to document collection operations.  Samples are carefully placed in suitable containers or coolers for prompt transportation to the laboratory.  Appendix D summarizes the type of sample container and preservation methods used, as well as the maximum acceptable holding time between sampling and analysis for various types of analyses and matrices.  When available and applicable, the holding times, sample container type and preservatives will follow regulatory guidance.

6.2.2    Sampling Site Identification

Sampling sites will be identified in a field logbook used for recording information during the conduct of sampling activities.  Each sampling site will be identified by exact location, which may include address, GPS coordinates, well number, or site name.  A unique sample site name and/or number is recorded in the field logbook and the sample collection form.  The sample site

name and/or number is also used to identify the sample on the record-of-custody form or chain-of-custody form.

    - Sample Container Inspection

Section No.     6

Revision No.     3

Date     03 Apr 00

Page      2 of 3

Inspect sample containers for good closure, proper labeling, and correct number and type required for the site.  Where split samples are being collected, additional containers will be needed.

6.3    Sample Collection

6.3.1	Sample Identification

When appropriate each sample will be uniquely identified by a number previously designated by the project/section officer.  This number will also be used on the project sampling forms.  The numbers assigned to splits, duplicate samples, and spiked samples will be coded in such a way to prevent easy identification as blind quality control samples when handled in the laboratory.  Labels with adhesive backings and with the sample number on the face will be

used.  Extra labels will be available.  Should more labels be required, they may be prepared with a permanent marking pen in the field, or a glass marking crayon may be used on the sample container.  In the latter case, an adhesive label should be prepared and attached to the sample container as soon as the sample is returned to the laboratory.  Ziplock bags used to carry samples will be labeled by writing appropriate identification directly on the bag using a permanent marking pen.

6.3.2    Collection of Field Replicate Quality Control Samples

Quality control criteria require that more than one set of samples be collected at a selected number of sampling events.  These samples will be used to verify the consistency of results.  Appropriate type and number of quality control samples will be specified with each project.

    - Field Storage of Samples

All sample containers will be kept in chilled storage in the field unless specific sampling protocol stipulates otherwise.  Insulated ice chests and frozen plastic-encased coolants (Blue Ice, for example) will be used.  For long term field storage of biological samples, dry ice will be used.  Ice will only be

used in sealed ziplock bags for transportation of live mussels to be transplanted.  The sampling team will have sufficient number of ice chests and frozen coolants to assure that samples remain chilled throughout the day.  The samples must always be kept in the possession of the sampling team until they are transferred to the custody of the laboratory.  Since the ice chests will have to be kept in a locked car or truck, the vehicle should be parked in the

shade to the extent possible.  Sampling vehicles should use unleaded fuels.  Ice chests will be cleaned with water and stored uncovered after each day.

Section No.     6

Revision No.     3

Date     03 Apr 00

Page      3 of 3

Sealed refrigerants will be washed with water and put into a freezer for reuse.  The vehicle will be refilled with fuel after samples are transferred when possible.

6.3.4    Storing and Shipping Samples

6.3.4.1  Storage at the Laboratory

The samples received at the laboratory will be kept in secured refrigerators or freezers.  Temperature will be kept as close as possible to the storage temperature required for each sample matrix and type of analysis.  Generally, refrigerators will be kept at 4 +/-2 degrees C, freezers will be kept at -15 +/-5 degrees C or colder.  Storage shall be in an environment where the sample identification numbers will remain attached.  Mechanical refrigeration units shall be used.  The use of ice as a refrigerant for sample storage is not allowed.

6.3.4.2  Shipping

All samples will be refrigerated or frozen during shipment through the use of cold packs or dry ice.  Ice is not acceptable as a refrigerant except for transportation of live mussels. Samples will be shipped in insulated containers.  All caps and lids will be checked for tightness prior to shipping.  To the extent possible, transporting vehicles will use unleaded fuel.

Section No.     7

Revision No.     2

Date     03 Apr 00

Page      1 of 3

7.0    SAMPLE CUSTODY

The Department of Fish and Game's chain-of-custody procedures for sample tracking are initiated during the time of actual sample collection by field personnel and maintained throughout the time the samples are in their possession.  Since the data generated will be used for either scientific or regulatory purposes, two custody procedures are necessary.  Chain of custody for regulatory samples that have potential for legal evidence, and Record of Custody for non-regulatory purposes.

7.1    Chain of Custody

The person responsible for sample collection must originate the chain-of-custody record. The sampler will clearly label the sample with the project name, sample location, field identification number, the date and time of sampling, and his/her own name and initials.  The same information will be entered on the chain-of-custody record along with information concerning the sample type, the analyses to be performed and the sample container. The individual collecting the samples will be responsible for the custody of samples until they are transferred or properly dispatched. If samples are hand-carried to the laboratory by Fish and Game personnel, custody of samples will be transferred to laboratory staff.   Samples transported by commercial carrier will have chain-of-custody seals affixed on the shipping container.  Documentation of the shipment will be kept with a copy of the chain-of-custody record by the person shipping the samples.  The original copy of the chain-of-custody record will accompany the sample(s) when transported to a departmental or commercial laboratory.

The laboratory staff person logging the sample(s) in will carefully inspect each sample for chain-of-custody documentation, sample labeling, packing lists, and for the condition of the custody seals, sample packing materials, and the sample containers.  Any discrepancies or problems associated with sample shipment will be documented on the chain-of-custody form.

After inspection, the samples will be entered into the laboratory sample receiving logbook, and will be assigned a unique sample identification number.  The following information shall be included when samples are logged-in:

- Laboratory number (assigned when samples are submitted)

    Laboratory storage location (refer or freezer no.)

    Spill Title (if applicable)

- Suspects name (if applicable)

- Index-PCA code (if applicable)

- Sampler’s name, address and phone number

- Date received by the laboratory

- Time and date sampled

Section No.     7

Revision No.     2

Date     03 Apr 00

Page      2 of 3

- Analysis requested

- Sample identification/location

- Sample type (matrix)

- Number of containers and container type

- Sample preservation

- Required report completion date

- Signatures of person submitting samples and person receiving samples for the

laboratory

- Date samples received by the laboratory

- Problem description (if applicable)

- Incident location (if applicable)

- Special instructions (if applicable)

The person logging the samples in will ensure that the samples are either retained in

secure storage or are given directly to an authorized analyst.  A copy of the chain-of-

custody form will be used to provide analysis requirements to the analyst(s).  This

form will accompany the sample containers and/or prepared extracts as each

authorized employee performs a required task on the samples.

After all analyses have been completed and disposal of the sample is authorized, a

designated sample custodian will make proper disposition of the sample with

appropriate documentation.  Disposal method and approximate disposal date will be

noted in the laboratory log-in records.  The completed chain-of-custody form(s) will

be retained as a permanent part of the project record.

  - Sample Handling, Storage, and Holding Times

All samples will be handled, prepared, transported, and stored in a manner designed

to minimize bulk loss, analyte loss, contamination or biological degradation.  The

sample containers will be clearly labeled with an indelible marker.  Soil and tissue samples for organic constituents must be frozen to prevent degradation or

volatilization.

Samples will be stored for the maximum sample holding time for the required

analyses as specified for the analysis.  Samples which do not have a maximum

holding time specified, will be stored for the duration of the research or study

activity unless the sample is consumed entirely for analysis.  Thereafter, the

laboratory supervisor or project leader will determine if the sample will be archived.

When the holding time interval has passed and samples are approved for disposal,

samples and the sample containers will be disposed of properly.  It is the sole

responsibility of the laboratory personnel to ensure that all applicable regulations

Section No.     7

Revision No.     2

Date     03 Apr 00

Page      3 of 3

are followed in the disposal of samples or related chemicals.  If the contracting officer

should request return of a sample prior to the maximum holding time, it will be

returned in a manner that meets Department of Transportation regulations.

Section No.     8

Revision No.     2

Date     12 Oct 92

Page      1 of 1

8.0    CALIBRATION PROCEDURES AND FREQUENCY

All laboratory instruments and equipment that are used for laboratory measurements, will be maintained and calibrated for good operating conditions that meet laboratory accuracy requirements.

The calibration/maintenance techniques will be performed according to a specific calibration standard operating procedure (SOP) which has been specified by the manufacturer's recommendations, an analytical or agency requirement, or by good laboratory practices.

Analytical standards to be used for instrument calibration are obtained from sources that have demonstrated accuracy levels, and are properly stored to ensure accuracy integrity.  Calibration chemicals will be logged in and assigned a specific shelf life based on chemical stability.  Non-chemical standards, instruments and equipment used for calibration purposes will be re-certified per an established schedule.

Routine maintenance will be performed by qualified laboratory personnel with recommended parts and supplies kept in stock.  For some items, the services of an outside vendor will be used.  The correct operation of instruments/equipment repaired by vendor services will be verified prior to use on projects.

An established schedule for the routine calibration or maintenance of the instrumentation and equipment will be developed based on manufacturer's recommendations, operational experience, procedural requirements and good laboratory practices. A maintenance log-book will be established for each instrument in which all calibration and maintenance techniques will be recorded.  Calibration results, which serve as a measure of instrument condition, must also be recorded in this manual.

The assigned laboratory quality control coordinator will make regular inspections of the log-books to ensure that calibration and maintenance is being performed at the required frequency.  If schedule times are exceeded, corrective action will be taken to substantiate data accuracy. Corrective action will also be initiated if the calibration results do not meet the required laboratory acceptance limits.

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      1 of 8

9.0    ANALYTICAL PROCEDURES AND DEVELOPMENT

For analysis mandated by regulatory agencies, specific methods have been designated for routine analyses with a  particular range of concentrations and matrices.  If methods are not stipulated, standard methods from a recognized authoritative source will be used for the tests.  Analytical methods and  reference sources routinely used at the Department of Fish and Game are listed in Appendix F.

9.1    Standard Procedures

All procedures developed and routinely used by the laboratory are documented with laboratory standard operating procedure (SOPs).  In addition to the actual test procedure, SOPs include applicable references, acceptance limits, health and safety precautions, trouble-shooting guidelines, quality control requirements, sample preparation and documentation criteria, calculation methods and reporting protocol. Analysts are trained to perform sampling and testing tasks per the established SOP.  Each analyst has access to a current copy of the procedure for the methods they perform.  When it is not possible to perform an analysis per established procedures, the area supervisor is promptly notified and corrective action may be initiated.

9.2    Method Development

If a suitable standard method is not available, the laboratory technical staff will develop appropriate methods to meet the project requirements.  Procedures developed within the laboratory are thoroughly validated to assure accurate, consistent results.  Reference materials, replicate analyses, and matrix spikes are some of the techniques applied to validate procedure development.  Based on these techniques, project-specific acceptance limits can be developed, and SOPs can be written.

Issued SOPs can be revised either on the basis of a periodic review schedule or an identified need.  The revisions are coordinated through the laboratory Quality Assurance Officer, and are distributed to appropriate personnel. A historical file is maintained on all SOPs and revisions.

  - Analytical Methods

Analytical methodology described in one of the following approved methodology manuals will be used as guidelines for all analytical methods used:

- EPA Methods for the Chemical Analysis of Water and Wastes, EPA-600/4-79-

020

- EPA Test Methods for the Evaluation of Solid Waste, Physical/Chemical

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      2 of 8

Methods, SW-846, third edition, 1986

- EPA Test Methods for Chemical Analysis of Municipal and Industrial

Wastewater, EPA-600/4-82-057

- EPA EMAP - Estuaries QAPP, EPA/600/x-93/xxx, May 1993

- Standard Methods for the Examination of Water and Wastewater,18th Ed., 1992

- Manual for Association of Analytical Chemists, 15th Ed.,1990

- USFWS, Patuxent Wildlife Research Center Analytical Manual (PWRCAM)

- Pesticide Analytical Manual (PAM Vol. I and II), USFDA

- Quality Assurance of Chemical Measurements by Dr. John Keenan Taylor

- Manual of Analytical Methods for the Analysis of Pesticides in Humans and

Environmental Samples, EPA-600/8-80-038.

Modifications of the above approved methods together with methodology developed by DFG personnel will be documented in SOPs.  This methodology may be used with approval in advance, in writing, by the Contracting Officer or the Contracting Officer's Technical Representative.  The limit of quantitation of any method which is used must meet a Method Detection Limit (MDL) consistent with the methodology being used. Analytical control will be maintained by strictly following written SOPs. If for some reason the SOP cannot be followed, deviations will be noted and reported in the data submital package.  Deviations from the approved analytical methodology must meet requirements based on the method validation test outlined subsequently in this section.

9.4    Analytical Method Validation

9.4.1    Limit of Detection

Determine, for each method, the limit of detection (LOD), defined as the lowest concentration level that can be determined to be statistically different from a blank, and the limit of quantization (LOQ), defined as the level above which quantitative results may be obtained with a specified degree of confidence.  Calculate the MDL and RL, the Federal Register, Vol. 49, No. 209, Friday, October 26, 1984 (Appendix G).

    - Standard Reference Material Test

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      3 of 8

To the extent that a standard reference material (NIST or National Research Council of Canada) can be obtained and when appropriate or required, it will be analyzed with each set of samples or every twenty samples for sets greater than twenty.  To validate a new or non-standard method for limited use, comparability data will be generated.

All results must be within 70-130% of the 95% certified confidence interval for the reference materials unless otherwise stated.

9.4.3    Sample Duplicate Test

For each set of samples, analyze one sample in duplicate for each analyte in question or one duplicate for every twenty samples.

    - Spike Recovery Test

One spike recovery test will be run for each set of twenty samples for each analyte on each matrix type analyzed.  For method validation, each matrix will be fortified with analyte at the reporting limit.  Recovery values for sample spikes must fall within 80 percent to 120 percent for method validation unless otherwise stated.

    - Method Documentation

Maintain a file containing all validation and modification reports. Upon completion of the validation tests, prepare a data report detailing the results.

9.5    Control Charts

Quality Control charts will be constructed as appropriate and maintained by each analyst to demonstrate that the laboratory is in a state of control.  Results of all quality control samples will be recorded on the appropriate quality control chart to determine the state of control.

9.6    Round-Robin Studies

The DFG laboratories will participate in round-robin studies and/or other standard reference sample programs as an on-going laboratory QC effort.

  - Organization of Laboratory

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      4 of 8

Qualifications of personnel is acknowledged to be very important to the laboratory.  When hired, personnel must have thorough knowledge of laboratory protocol and

suitable experience.  Such experience can be obtained from laboratory experience in another facility or suitable college class work.  Anyone conducting analytical procedures in the laboratory is responsible for the accuracy of those procedures and is answerable to the Laboratory Director.  New personnel will be trained by a qualified analyst and will report to the appropriate Project/Section Leader.  All personnel are expected to be familiar with and carefully follow the  appropriate laboratory standard operating procedures (SOPs) developed for use in the laboratory when conducting analyses on samples received.

9.8    Laboratory Operating Practices

9.8.1    Sample Receipt

The following procedure will be followed immediately upon receiving a sample shipment:

•	    Record sample number in sample log book and check that the sample number is clearly marked on the sample container and on accompanying custody forms or sample worksheets.

•    Record the requested analytical information in the sample log book.

•    If sample is not going to be run immediately, follow appropriate sample

storage procedures.

9.8.2	Laboratory Procedures

The individual analyst upon receiving a sample shipment will proceed as

follows:

•	Record date and procedure to be used.

•	Record all pertinent information regarding instrument and/or materials

in the lab book or on the data sheet.

•	Examine glassware routinely to confirm cleanliness.

•	Check to make sure that the reagents used are the correct grade and

type for the analysis to be done.

•	If a new lot of reagent is put into use, reagent blanks or other checks

should be run to demonstrate continuity of the required quality.

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      5 of 8

•	Follow designated procedure (SOP) exactly, including all quality

control requirements.

    - Record Keeping

The following logs will be maintained by laboratory personnel:

•    Sample log: records date of receipt, number of samples, laboratory

sample number, analyses to be completed.

•	Run log:  records each analytical run by run number, date, and

analysis.  Data printouts shall be referenced Laboratory number (log-in

or L#).

•	Standards preparation log:  records all weights and volumes  used to prepare standards, solvent, source and purity or concentration of neat or concentrated standards, date prepared, final concentration, and preparer's initials.  All standard storage containers will be labeled to reference the standard preparation log.

•    	Lab book/worksheet:  record information pertinent to sample preparation and method used, instrument calibration and parameters, raw data and calculations, and any other information deemed important to the documentation of analyses done.

    - Reports

Procedures for analytical reporting will be as follows:

•    	Retain all appropriate computer printouts and strip chart recordings in

a binder (be sure information includes sample number, date and time).

•    Have all information clearly recorded so that a written data report can

be made and reviewed.

•    Individual data reports shall be given to the Laboratory Supervisor for

compilation and review prior to the preparation of the final report.

    - Instruments

Instrument check procedures will be implemented as follows:

•    The instrument demonstrates linearity or is calibrated for the range in

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      6 of 8

which work is intended or the instrument program or parameters being used will be verified.

•    The accessory equipment (such as syringes or autosampler tubes)

must be the correct ones for the instrument and method being used

and they must be CLEAN.

•    The reagents being used are the proper ones, and they have

been checked for interferences by running a blank.

•    When beginning a procedure, list the page number of calibration

run and blanks in log book.

•    When results are obtained, record printout page number in log book.

•    If the instrument or other equipment does not appear to be working properly, contact the supervisor IMMEDIATELY.

    - Quality Assurance Procedure

The following quality assurance procedures form the foundation for quality control practices in the laboratory.  These procedures will be practiced as a matter of routine unless superceded or modified by specific quality assurance requirements of a given project.

      - Chemical Analysis – General

•    Run blanks of all new reagents used - record date, time, and

page number of printout in log book.

•    At least one method blank will be run for each set of samples

of one matrix type to determine whether interferences are introduced.  Method blanks will be run through the complete analytical method along with the sample set.  Blanks shall be run with a minimum frequency of one blank per 20 samples unless otherwise specified.

•    At least one sample will be fortified or one SRM of similar

matrix to that of the samples will be used and run with each

set of samples to measure the analytical accuracy.  When

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      7 of 8

appropriate the sample will be spiked with the analyte(s) at  the

expected level in the sample or at mid- range of the standard

curve.  A minimum of one fortified sample or SRM will be run

with each sample set of 20 samples or less.

•    At least one sample will be prepared in duplicate and analyzed

with each set of samples of one matrix type to measure

analytical precision.  A minimum of one duplicate will be

analyzed for each set of 20 samples or less.

•    Prepare new standards as needed according to the guidelines in

section 8.0.  Record all the required information in the

standards preparation log.  Freshly-made standards should be

compared to the response of existing calibration standards

before using in order to verify the reliability of the new

standard.

•    When washing glassware, follow procedure outlined in Standard

Operating Procedure (Appendix C).

•    Standard deviation calculations will be performed routinely at

the end of each run.

•	    All QA procedures used will be listed in the log book or SOP.

•    Additional periodic checks of accuracy or precision will be

required as deemed necessary by laboratory supervisor.

      - Review of Quality Control/Analytical Data

•    All quality control/analytical data will be reviewed for

correctness of the analytical, calibration, and data reduction

procedures used, and initialed by the section leader or

laboratory supervisor before the accompanying data may be

reported.

•    If after being reviewed, a set of data are determined to be out of

control, the quality assurance coordinator shall be notified and

an appropriate course of corrective action will be prescribed.

(See sections 11, 14, and 15 for data evaluation criteria and

corrective measures.)  The analyst shall enter the corrective

Section No.     9

Revision No.     4

Date     03 Apr 00

Page      8 of 8

measures taken in the lab notebook, which will then be signed

by the supervisor or QA officer.  No additional  analytical data

will be generated until the problem has been identified and corrected.

    - Laboratory Safety

•	Safety glasses will worn in designated laboratory areas at all times.

Laboratory visitors will be issued safety glasses when necessary.

•	Lab coats or aprons will be worn while working with any solvents,

acids, caustics, condensers, or designated instrumentation.

•	Walkways, exits, and safety shower/eyewash stations will be clear of

debris at all times.

•	No horseplay of any kind will be tolerated.

•	No food or drink will be allowed in laboratory areas.

•	Additional safety and emergency procedures outlined in the Standard

Operating Procedures will be followed.

•	Refer to laboratory safety manual for additional safety measures.

    - Laboratory Cleanliness

•	Glassware will be washed and prepared for use in accordance with the

procedures set forth in Appendix C (Standard Operating Procedures).

•	Clean glassware will be returned to its proper place as soon as it has

been cleaned appropriately and properly capped.

•	Counter tops will be kept clean.

•	Spills will be cleaned up immediately.

•	Broken glassware will be placed in a specified container.

9.8.9   Records

Records on all relevant data are to be easily located in files that pertain to a specific analysis, question, or project.

Section No.     10

Revision No.     3

Date     03 Apr 00

Page      1 of 2

10.0    DATA REDUCTION, VALIDATION AND REPORTING

  - Analysis

The conversion of raw data into functional values and the presentation of these

values is a critical process in the laboratory function.  In order to assure the

production of data that is scientifically valid, defensible, comparable and of known

precision and accuracy, the following steps are required.

  - Validation

Reduction of raw data is accomplished using established techniques.  The

calculations required to perform the reduction of raw data are performed manually

or with the aid of automated data processing systems, as specified by the SOP for

the particular testing event.  If manual processing is to be used, the SOP will

provide the calculation method and the units for reporting derived values.  For

automated data reduction, the accuracy of calculations will be verified through the

use of standards or raw data inputs of known values.

Raw data, related quality control information and derived values are carefully

evaluated prior to final reporting.  The initial evaluation is performed by the

analyst/specialist performing the work.  Statistical methods, such as precision and

accuracy acceptance limits and/or control charts, are employed to assess data

acceptability.  The laboratory supervisor provides a second evaluation of the data

and conclusion contained in the final report.

  - Reporting

Analytical reporting limits will be experimentally developed either as instrument

detection limits or method detection limits (MDL).  For applications requiring a

greater degree of statistical confidence, the reporting limit (RL) will be used to

establish a minimum reporting limit.  The minimum reporting level applied, will be

based upon project requirements and proven laboratory capabilities.

  - Final Reports

The final reports contain an outline of the scope of the project, sample

identification, methodologies performed, a discussion of any unusual circumstances

regarding the project, and tabulated analytical results. This report is reviewed and

signed by the person who prepared the analyses, and by the technical reviewer or

supervisor.

  - Record Keeping and Maintenance

Section No.     10

Revision No.     3

Date     03 Apr 00

Page      2 of 2

Instrument log books will be maintained with each instrument.  All analytical parameters are to be recorded for each run and calibration data recorded daily.  In addition, a record will be made of any conditions or incidents the analyst encounters which are in any way unusual, or deviate from the SOP.

When maintenance is required, a record will be made of the symptom, the repair

performed, and the individual performing the repair.

All observations, tapes, and most printouts, and other raw material generated in

the course of any analysis will be saved.  They will be filed with reference to

laboratory log number, date, batch number, analyst, and other information deemed

pertinent.  Also recorded in the laboratory notebook will be all weights or other

types of raw data generated in the laboratory but not printed on a hard copy by the

data generating device.  All documentation in the laboratory notebook will be

made in ink.  Corrections to notebooks or other data records will be made by

crossing a single inked line through the error, entering and initialing the

correction, and recording the date.

The records to be maintained in the laboratory include such items as sample

tracking records, analysts' notebooks, bench sheets, instrument read-out records,

computer printouts, quality control data, and raw data. The records will be

maintained for the life of the project and they will be provided to the

organizations contracting for the research or studies upon request.

Section No.     11

Revision No.     3

Date     03 Apr 00

Page      1 of 3

- INTERNAL QUALITY CONTROL CHECKS AND FREQUENCY

Quality control checks are routinely performed in the Fish and Game laboratory operations.

These checks may be increased or modified to meet the needs of a particular analysis.

  - QA Samples

Internal Quality Assurance samples (appropriate reference materials) will be

analyzed with each set or every twenty analyses being performed.  These internal

quality assurance analyses are conducted for the parameters being monitored by that

analytical procedure.  In addition, the compounds contained in the quality assurance

sample will be representative of those compounds being monitored.  Accuracy

(Relative Percent Error) is calculated as follows:

Accuracy = [(Da-D)/D]x100   (Absolute Value)

Where, Accuracy = relative percent error

Da       = Analysis value of QA sample

D        = Accepted value of QA sample

The results of all QA analyses and the percent deviation from the accepted

analytical values will be recorded in the laboratory notebook.  For accuracy

(percent deviation from accepted value) use a control limit of 20 percent for sample

values greater than 30 times the MDL.  For concentrations less than twenty times

the accepted value times 100 will be used as the control limit.

  - Duplicate Samples

One or more duplicate samples will be analyzed for each set of twenty samples

analyzed.  The relative percent difference for each constituent is calculated as

follows:

RPD = {(D1-D2)/[(D1+D2)/2]}x100

Where,  RPD = Relative Percent Difference

D1      = First Sample Value

D2      = Second Sample Value (duplicate)

The results of all duplicate determinations and the calculated relative percent

difference will be recorded in the laboratory notebook.  For RPD, use a control limit

Section No.     11

Revision No.     3

Date     03 Apr 00

Page      2 of 3

of 15 percent for sample values greater than 30 times the MDL.  For concentrations

greater than 5 times and less than 30 times the MDL, a difference of 2 MDL or 20

percent will be permitted, whichever is greater.  For concentrations less than 5 times

the MDL, a difference of one MDL will be permitted.  If either sample value is less

than the MDL, the RL will not be calculated; in its place the notation of "ND" will

be recorded in the laboratory notebook.  If the precision falls outside the control

limits, the analyses must be terminated, the problems corrected, and the previous

samples associated with that duplicate reanalyzed.  If duplicate results are outside of

the control limits, all pertinent data for samples associated with that duplicate

sample will be flagged in the laboratory notebook.

  - Matrix Spiked Sample Analyses

When required matrix spike and matrix spike duplicate analyses will be conducted at a rate of five percent.  The spike will be added prior to any digestion, extraction, or distillation steps as a check on the sample preparation and methodology.  An amount of analyte will be added to the sample that is ten times the MDL or equivalent to the analyte concentration.  Recovery values are calculated as follows:

Recovery = [(Ds-D)/S]x100

Where, Recovery = Relative Percent Recovery

Ds       = Value of Sample with Spike

D        = Value of Sample without Spike

S        = Amount of Spike Added

Recovery values for sample spikes must fall within 80 percent to 120 percent except

where a specific method and analyte require a different acceptable range.

Exceptions shall be noted in the method protocol.  When a specific method and

analyte require a different acceptable recovery range, as determined by actual spike

recovery runs, the acceptable range shall be noted in the Standard Operating

Procedure for that method.  If the recovery falls outside of the acceptable recovery

range, the analyses must be terminated, the problems corrected, and the previous

samples associated with the spike reanalyzed.  When sample concentrations are less

than the MDL, the value of "0" will be used as the sample result concentration for

purposes of calculating spike recoveries.  All spiked sample results will be recorded

in the laboratory log book and reported.

  - Method Blanks

Method blanks will be analyzed at a minimum of once for every batch of samples or

at a rate of five percent of the samples analyzed.  Blank values must not exceed

three MDL.  If blank values exceed three MDL, the analysis must be terminated, the

Section No.     11

Revision No.     3

Date     03 Apr 00

Page      3 of 3

problems corrected, and the previous samples associated with the blank re-analyzed.

All blank analysis results will be recorded in the laboratory log book and reported.

Section No.     12

Revision No.     4

Date     13 Jul 01

Page      1 of 1

- PERFORMANCE AND SYSTEM AUDITS AND FREQUENCY

The two types of audits implemented in Office of Spill Prevention and Response (OSPR)are the system audit and the performance audit.  The system audit is an on-site review which provides a qualitative appraisal of a department's quality assurance system and facilities for sampling and measurement.  The performance audit qualitatively or quantitatively assesses adherence to laboratory SOPs and required quality work performed by technical personnel.

  - Performance Audit

The performance audit system consists of an employee performance audit by the

laboratory supervisor and senior technical personnel.  This audit is implemented

at least once  yearly to verify satisfactory performance.  If acceptable results are

not obtained for the performance audit program, corrective action will be initiated.

  - System Audit

The Quality Assurance Officer or person acting in that capacity conducts a

QA/QC evaluation of each project when data are reported.  This evaluation

includes a review of QC data, log books and other documentation.  A report of the

findings is submitted to the project officer.  The project officer may be required to

submit a corrective action plan for all unsatisfactory findings on an evaluation.  In

addition, the QA/QC Officer audits all operations on an on-going basis, for conformance to the required procedures.

Section No.     13

Revision No.     1

Date     14 Mar 90

Page      1 of 1

- PREVENTATIVE MAINTENANCE

Maintenance on analytical instruments will be performed by in- house instrument maintenance personnel or by manufacturers service personnel.  Critical spare parts (for gas chromatographs - septa, syringes, column ferrules, backup column, etc.; for atomic adsorption spectrophotometers -- AA lamps, quartz cell, plastic tubing, etc.) will be maintained on hand.

Fume hoods will be checked quarterly.  The results and the ventilation capacity across the face of the fume hood and the inspection date will be posted on an exterior wall of the fume hoods.

Equipment manuals containing trouble-shooting SOPs will be kept either with the instrument or in the possession of in-house maintenance personnel.

Instrument operators are responsible for daily maintenance and for maintaining instrument logs.  These logs will contain the date and description of routine maintenance procedures.  Each entry into the run log will be initialed by the individual making the entry.  The run log will be reviewed periodically by the Quality Assurance Officer.

Section No.     14

Revision No.     1

Date     14 Mar 90

Page      1 of 2

- ROUTINE ASSESSMENT OF DATA PRECISION, ACCURACY, AND

COMPLETENESS

  - Precision

Precision shall be assessed with each sample set for each analysis type. Precision

will be expressed in terms of relative error as the percent deviation of the

duplicate results from the original results obtained.  The equation for determining

precision is:

RPD = (D1-D2)/[(D1+D2)/2]x100

Where  RPD = Relative Percent Difference

D1    = First Sample Value

D2    = Second Sample Value (duplicate)

  - Accuracy

Accuracy will be assessed on a regular basis in each set of samples for each

analysis type by comparison of the analytical results of internal QA samples

provided or approved by the QA Officer with accepted concentrations.  Accuracy

will be expressed in terms of relative error as the percent deviation of the

analytical results from the known values.  Accuracy is calculated as follows:

Accuracy = [(Da-D)/D]x100

Where, Da = Analysis value of QA sample

D  = Accepted value of QA sample

In the absence of internal QA samples, accuracy will be assessed using spiked

samples.  Recovery of the spike will be used in this instance. Recovery is

calculated as follows:

Recovery = [(Ds-D)/S]x100

Where, Recovery = Percent Relative Recovery

Ds       = Value of Sample with Spike

D        = Value of Sample without Spike

S        = Amount of Spike Added

  - Completeness

Section No.     14

Revision No.     1

Date     14 Mar 90

Page      2 of 2

Completeness shall be assessed for each sample set and for each analysis type.

The comparison for completeness will consist of a comparison of the number and

type of analyses scheduled to be performed with those analyses successfully

completed.  Completeness shall be expressed as the percentage of analyses

successfully completed relative to the number of analyses scheduled to be

performed for each analysis type.

Section No.     15

Revision No.     2

Date     12 Oct 92

Page      1 of 1

- CORRECTIVE ACTION

Corrective action includes a variety of activities starting with the individual analyst applying the elements of quality control to a particular task.  At this level, the corrective action takes the form of problem identification based on spike, calibration, or recovery results that exceed acceptance limits.  Appropriate action to be taken at this stage includes checking calculations or calibrations, preparing new standards or spiking solutions, or re-analyzing samples.  If the above actions do not correct the matter, the laboratory supervisor is notified. Project work requiring the use of the defective instrument will be suspended until the problem has been resolved.

System audits, reference standards, blind duplicates and standards may also indicate a necessity for corrective action.  In these instances, corrective action for out-of-limit values are normally requested by the laboratory Quality Assurance Coordinator.

The data generated during the period of problem identification will not be reported, unless additional analysis is not possible due to restricted sample availability, or time constraints.  In this case, the result will be reported with qualifications that have been clearly identified and approved by project officials.

Section No.     16

Revision No.     2

Date     03 Apr 00

Page      1 of 1

- QUALITY ASSURANCE COMMUNICATION WITH MANAGEMENT

Communication with all levels of management concerning quality subjects is an ongoing process and routine quality issues are communicated to appropriate levels of management.  The results of performance audit evaluation sample analyses are routinely provided to the laboratory supervisor, as well as to the project officers.  If significant QA problems are experienced or observed in any aspect of lab operations, the QA officer and management are promptly notified along with the recommended solution.

The results of all system audits are communicated to designated levels of management.  The findings of system evaluations conducted by the division's quality assurance/quality control officer are provided both to the laboratory supervisor and the project officers.

Section No.     17

Revision No.     0

Date     14 Apr 00

Page      1 of 2

17.0	STAFF TRAINING AND DOCUMENTATION

  - Hiring Process

People often begin their careers at CDFG as temporary employees.  Under

these circumstances, screening and reference confirmation is undertaken by

the Department, and copies of that information are retained in the person’s

newly created employee folder.  If/when the employee becomes

permanently hired by CDFG, she/he signs the agency information, and those

papers are returned to the personnel file.

New employees hired by CDFG provide documentation of skills they

already possess via publications, detailed resumes, letters of

recommendation, and self-certification.  Subsequent staff training is

documented via training checklists and written management and/or peer

reviews, which are maintained in chronological order in the personnel files.

Every staff member is formally evaluated annually using a combination of

self and supervisor evaluation.  Job descriptions are reviewed and may be

updated at that time.

  - Training

Employees trained in a new skill learn by a mentorship process.  The

employee is assigned to a staff member, who teaches each method as

follows:

a)	First familiarization with the equipment/process by observation

b)	Detailed reading of the SOP and attendant literature references

c)	Second familiarization with the equipment/process by observation

d)	Supervised practice of the method using standards and SRM’s

e)	Unsupervised practice of the method on standards and SRM’s

f)	Blind re-analysis of actual project samples, which have also been

independently analyzed/processed by the mentor

g)	Evaluation of the blind intercomparison by the new staff member,

mentor, and laboratory manager

- Supervised analysis of low QC level samples or samples generated by internal research projects.

Completion of these steps is documented using a training checklist, signed by the trainee and by her/his supervisor.  Once the new employee has successfully inter-compared on at least ten samples (i.e., all results are found to replicate the mentor’s results within ± 20%) she/he is considered trained in that task, and a note is placed in her/his personnel file.

Section No.     17

Revision No.     0

Date     14 Apr 00

Page      2 of 2

For newly trained employees, pertinent QC data is reviewed on a regular basis by the laboratory director or project lead chemist for a period of at least one week, until the lead chemist and/or laboratory director is satisfied that the employee is competent in the procedure.

As the employee sets up her/his equipment and begins to obtain actual data, ongoing performance evaluations occur via control chart information, inter-laboratory inter-comparison exercises, and blind Performance Evaluation (PE)samples documenting precision and accuracy.  Records of performance are submitted by the staff member and/or lead chemist (quarterly basis) to the lab director, for inclusion in the staff member’s personnel file.

Training documentation differs according to employment category.  Principle Investigator’s training is primarily documented by the quantity and quality of complete laboratory reports.  For technical staff, CDFG relies on documented analytical performance criteria (accuracy and precision on blind samples).  Copies of all documentation are maintained

by the QA officer (acting).

  - Seminars and Meetings

Records are also maintained on all seminars, classes, and training sessions attended by each employee. On a yearly basis, all staff attend at least one safety seminar and one QA/QC review seminar, where current practices and new procedures are discussed and overall laboratory compliance is reviewed.  In addition, staff attend the laboratory’s regular staff meetings.  The concept of the meetings varies and speakers may be invited from outside sources or may be CDFG staff.

In the event that any serious safety or QA/QC deficiencies are discovered by the lab director, or if a new employee is hired, or if new project-specific requirements are mandated, a staff meeting will immediately be called to provide the appropriate information and establish any required protocols.  The agendas from all staff meetings which include training issues, health and safety or discussion of QA/QC concerns will be signed by the attendees, and copies will be retained in personnel files.

Section No.     18

Revision No.     1

Date     12 Mar 92

Page      1 of 2

18.0  Glossary

Accuracy - combination of bias and precision of an analytical procedure, which

reflects the closeness of a measured value to a true value.

Bias - consistent deviation of measured values from the true value, caused by

systematic errors in a procedure.

Calibration check standard - standard used to determine the state of calibration of an

instrument between periodic recalibrations.

Confidence coefficient - the probability, %, that a measurement result will lie within

the confidence interval or between the confidence limits.

Confidence interval - set of possible values within which the true value will lie with a

specified level of probability.

Confidence limit - one of the boundary values defining the confidence interval.

Detection limits - Various limits in increasing order are:

Instrument detection limit (IDL)-the constituent concentration that produces a signal greater than five times the signal/noise ratio of the instrument.  This is similar in many respects, to "critical level" and "criterion of detection." The latter limit is stated as 1.645 times the s of blank analyses.

Lower limit of detection (LLD) - the constituent concentration in reagent water that

produces a signal 2(1.645) s above the mean of blank analyses.  This sets both Type I and Type II errors at 5 %.  Other names for this limit are "detection limit" and

"limit of detection" (LOD).

Method detection limit (MDL) - the constituent concentration 	that, when processed

through the complete method, produces a signal with a 99% probability that it is

different from the blank. For seven replicates of the sample, the mean must be

3.14s above the blank where it is the standard deviation of the seven replicates.

The MDL will be larger than the LLD because of the few replications and

the sample processing steps and  may vary with constituent and matrix.

Limit of quantization (LOQ) - the constituent concentration that produces a signal

sufficiently greater than the blank that it can be detected within specified limits by

good laboratories during routine operating conditions.  Typically it is the  concentration that produces a signal 10s above the reagent water blank.

Duplicate - usually the smallest number of replicates (two) but specifically herein

refers to duplicate samples, i.e. two samples taken at the same time from one

location.

Internal standard - a pure compound added to a sample extract just before

instrumental analysis to permit correction for inefficiencies.

Laboratory control standard - a standard, usually certified by an outside agency, used

to measure the bias in a procedure.  For certain constituents and matrices, use

National Institute of Standards and Technology (NIST)* Standard Reference

Materials when they are available.

Precision - measure of the degree of agreement among replicate analyses of a sample,

usually expressed as the standard deviation.

Section No.     18

Revision No.     1

Date     12 Mar 92

Page      2 of 2

Quality assessment - procedure for determining the quality of laboratory

measurements by use of data from internal and external quality control measures.

Quality assurance - a definitive plan for laboratory operation that specifies the

measures used to produce data of known precision and bias.

Quality control - set of measures within a sample analysis methodology to assure that

the process is in control.

Random error - the deviation in any step in an analytical procedure that can be treated

by standard statistical techniques.

Replicate - repeated operation occurring within an analytical procedure.

Surrogate standard - a pure compound added to a sample in the laboratory just before

processing so that the overall efficiency of a 	method can be determined.

Type I error - also called alpha error, is the probability of deciding a constituent is present when it actually is absent.

Type II error - also called beta error, is the probability of not detecting a constituent when it actually is present.

-----------

*Formerly National Bureau of Standard (NBS).

APPENDIX A

ORGANIZATIONAL CHART

APPENDIX B

QUALIFICATIONS AND SPECIFICATIONS OF KEY PERSONNEL

(Resumes Available upon Request)

SPECIFICATION FOR KEY PERSONNEL

Specifications for QC Coordinator

1.  Respected person, have authority

2.  Laboratory experience 5 - 10 years

3.  Supervisory experience 5 - 10 years

4.  Safety committee candidate

Specifications for Organic Laboratory Data Validation Project Leader

1.  Laboratory experience 5 - 10 years

2.  Gas chromatography experience 2 - 3 years

3.  Mass spectral interpretation experience 2 - 3 years

4.  Communication skills

5.  Computer skills

6.  Dedicated to improvement

Specifications for Inorganic Laboratory Data Validation Project Leader

1.  Laboratory experience 3 - 5 years

2.  Atomic absorption  experience 2 years

3.  Communication skills

4.  Computer skills

5.  Dedicated to improvement

Specifications for Laboratory Personnel

1.  Bachelors degree in chemistry or related sciences

2.  Organic extraction experience 1 year

3.  Inorganic digestion experience 6 mo.

4.  Synthetic organic residue experience 2 years

APPENDIX C

SOP’s

(Standard Operating Procedures for specific methods are available on request)

APPENDIX D

SAMPLE CONTAINERS, PRESERVATION AND HOLDING TIME

Samples Analyzed for Synthetic Organics

Polynuclear Aromatic Hydrocarbons or Petroleum Hydrocarbons

| Sample Type | Sample Size | Containers | Preservation | Holding Time |
| --- | --- | --- | --- | --- |
| Water | One gallon | Glass 1,2 | 4o C, pH 5-9 | 7 days 3 |
| Animal | Whole | Al foil | -20o C | 6 mo. |
| Vegetation | One pint | Al foil | -20o C | 6 mo. |
| Sediment | One pint | Glass 1 | -20o C | 14 days 2 |

1.  Previously rinsed with petroleum ether and dried, with Teflon liner in lids.

2.  Sample must be extracted within the specified days and analyzed within 40 days of extraction.

3.  PAHs are light sensitive, therefore, sample extracts and standards must be stored in foil

wrapped containers.

Samples Analyzed for Trace Elements

| Sample Type | Sample Size | Containers | Preservation | Holding Time |
| --- | --- | --- | --- | --- |
| Water | 500 ml | LPE1 | HNO3 to pH<2 | 6 mo.2 |
| Animal | Whole | Plastic bag | -20o C | 6 mo. |
| Sediment | One pint | LPE1 1 | -20o C | 6 mo. |

1.  Previously soaked and rinsed with 1N HNO3.

2.  Six months except mercury and TBT which are 28 days.

APPENDIX E

FORMS

APPENDIX F

ANALYTICAL METHODS AND REFERENCE SOURCES

APPENDIX G

METHOD DETECTION LIMIT AND REPORTING LIMIT