- OMI/Thames Water Stockton, Inc

- LABORATORY

- QUALITY ASSURANCE MANUAL

- Revision 4.20

03/02/2007

_______________________________________________________________                                                                  ____________________________________

Larry Huber							Date

Laboratory Supervisor

- Table of Contents

Introduction	1

1.  Laboratory Description	2

2.  Laboratory Organization and Responsibilities	2

2.1  Laboratory Supervisor	2

2.2  Chemist	2

2.3  Microbiologist	3

2.4  Laboratory Technician	3

3.  Quality Assurance Procedures and Objectives	3

3.1  Objective	3

3.2  Procedures	3

3.2.1  Methods	3

3.2.2  Training	3

3.2.3  Evaluation	3

3.2.4 Corrective actions	3

4.  Sampling Procedures	3

4.1  Sample Collection	3

4.2  Sample Containers	4

4.2.1  Container Material	4

4.2.2 Container Preparation	4

4.2.2.1  General procedure for all containers	4

4.2.2.2  Containers for Metals Analyses	4

4.2.2.3  Containers for Bacteriological Samples	4

4.2.2.4  Containers for Organic Samples	5

4.3  Sample Preservation and Storage	5

4.3.1  Required monitoring samples	5

4.3.2  Other samples	5

5.  Sample Custody	5

5.1  Sample Custody	5

5.2 Routine Samples	5

5.3  Special Samples	6

5.3.1  In-house Chain of Custody	6

5.3.2  Transfer of Custody	6

6.  Calibration Procedures and Frequency	7

6.1  Meters	7

6.2  Calibration Curves	7

6.3  Personnel	7

7.  Analytical Procedures	7

7.1  References	7

7.2  Methods, Containers, Preservation, and Holding Times	8

8.  Data Reduction, Validation, and Reporting	11

8.1  Data Reduction	11

8.2  Data Validation	11

8.2.1  Quality assurance review	11

8.2.2  Final Approval	11

8.3 Reporting of results	11

9.  Internal Quality Control Checks	11

9.1  Types and purposes of method QC procedures	11

9.1.1  Method (Reagent) Blank	11

9.1.2  Duplicate (Replicate) Sample	11

9.1.3  Method Spike	12

9.1.4  Sample Spike	12

9.1.5  Reference Sample	12

9.1.6  Check Sample	12

9.1.7  Reporting Limit Verification Sample	12

9.2  Types and purposes of sample QC procedures	12

9.2.1  Container Blank	12

9.2.2  Trip Blank	12

9.2.3  Field Blank	13

9.2.4  Field Duplicate	13

9.3  Analytical Reagents	13

9.3.1  Chemicals	13

9.3.1  Purified water	13

9.3.2  Distilled water	13

9.4  Standards	13

9.4.1  Metal and fluoride standards	13

9.4.2  pH Buffers	13

9.4.3  Other standards	13

9.5  Methods	13

9.6  Recording of data	13

9.7  Calculation and reporting of analytical results	14

9.8  Review of analytical results	14

9.9  Acceptable limits of analytical results	14

9.10  Definitions of analytical limits	16

9.10.1  Method Detection Limit (MDL)	16

9.10.2  Instrument Detection Limit (IDL)	16

9.10.3  Reporting Limit (RL)	16

9.10.4  Minimum Level (ML)	16

9.10.5  Spike Recovery (%S)	16

9.10.6  Percent Range (%R)	16

9.11  General Requirements	16

9.11.1  Reagent Blank	16

9.11.2  Standards	16

9.11.3  Sample controls	16

9.11.4  Performance/maintenance records	16

9.11.5  Data	16

9.11.6  Calculations	16

9.11.7  Non-routine samples	17

9.12  Bacteriology	17

9.13  Biotoxicology	17

10.  Performance and System Audits	17

11.  Preventive Maintenance	17

11.1  General laboratory equipment	18

11.2  Maintenance Contracts	18

12.  Assessment of Precision and Accuracy	18

12.1  Precision	18

12.2  Accuracy	18

12.2.1 Checks	18

12.2.2  Reference samples	18

13.  Corrective Actions	18

14.  Quality Assurance Reports	19

Introduction

This document provides a description of the minimum monitoring and control requirements intended to insure the reliability of results reported by this Laboratory.

Such requirements include adherence to approved methods by trained personnel, adequately maintained equipment, quality control checks to monitor the accuracy and precision of results, and review and documentation of results and the means by which those results are obtained.

1.  Laboratory Description

The laboratory is an operating section of OMI/Thames Water Stockton, Inc (OMI/TW).  OMI/TW operates the Wastewater Division of the City of Stockton, Municipal Utilities Department under a service contract and is located at the City's Regional Wastewater Control Facility on Navy Drive in Stockton, California.  The laboratory facilities and staff are the same as before the transition to the service contract.  Analytical services are routinely provided to the Wastewater, Stormwater, and Water/Sewer Divisions.  Facilities consist of a main laboratory located in the City administration building at 2500 Navy Dr., a trailer permanently installed near the discharge point at the tertiary treatment plant for performing flow-through bioassays, and a boat for conducting receiving/ambient water monitoring and sample collection.  All testing except for bioassays and field measurements are performed in the main laboratory.  Services may be provided on an as needed basis to other departments within the City, and occasionally to outside agencies.

The scope of work performed by the Laboratory includes numerous sampling programs, and analyses for NPDES monitoring requirements, treatment plant process control, industrial wastes (Pretreatment program) and Title 22 monitoring requirements for public water supply systems.  The Laboratory has the capability of performing most of the required analyses, except trace metals, organic and radiochemical.  Analyses that cannot be performed in-house are contracted to certified commercial laboratories.

- 2.  Laboratory Organization and Responsibilities

2.1  Laboratory Supervisor

The Supervisor is responsible for the overall operation of the laboratory.  This includes, but is not limited to:

Personnel selection (with other management staff); upgrading facilities, methodologies and equipment; reviewing data and the performance of the Laboratory; establishing and maintaining of the Quality Assurance Program; directing the work of the Chemists, Microbiologist, and Laboratory Technicians; preparing monitoring reports for regulatory agencies.

The Laboratory Supervisor reports to the Technical Services Director.

  - Chemist (2)

A Chemist reports to the Laboratory Supervisor and may perform or oversee the performance of trace metal analyses or other more sophisticated equipment and methods; is a lead person for and directs and coordinates the work of several Laboratory Technicians; checks worksheets for accuracy and completeness; reviews and compiles data; maintains the Quality Assurance Program; supervises and/or develops laboratory methodologies.

- May assume and perform the duties of the Supervisor in his absence.

2.3  Microbiologist (1)

The Microbiologist reports to the Laboratory Supervisor and performs bacteriological and biotoxicological sampling and testing; is a lead person for and directs the work of one Laboratory Technician; reviews and compiles data; maintains the Quality Assurance Program; supervises and/or develops laboratory methodologies.

- May assume and perform the duties of the Supervisor in his absence.

2.4  Laboratory Technicians (5)

Laboratory Technicians are immediately responsible to their lead person but report ultimately to the Laboratory Supervisor; perform sampling and analyses according to established procedures; participate in the Quality Assurance Program.

- 3.  Quality Assurance Procedures and Objectives

3.1  Objective - The objective of the Quality Assurance Program is to assure the accuracy and precision, as well as reliability of laboratory results.  Emphasis will be on the prevention of quality problems rather than detection and correction of problems after they occur.

3.2  Procedures - Procedures that will be used to accomplish this objective include:

3.2.1  Methods will be chosen which have been proven and accepted by regulatory agencies, and which have been shown to be reliable and provide sufficient accuracy and precision.

3.2.2  Training - Staff members will be trained in using the analytical methods, equipment, and the basic concepts of quality control to make them effective in supporting quality assurance efforts.

3.2.3  Evaluation  - Overall performance of the laboratory and adherence to quality policies will be evaluated periodically to assess the quality level of the work, and to determine where changes are necessary.

3.2.4  Corrective actions will be taken to cure problems and to strengthen areas where increasing quality controls could prevent future problems.

- 4.  Sampling Procedures

4.1  Sample Collection

- Samples will be collected in appropriate containers and will be representative of the source being sampled.

- The field personnel will insure proper labeling and are responsible for the integrity of the sample until it is returned to the Laboratory.

- The Laboratory Supervisor will determine whether proper custody procedures were followed during the field work and will decide if additional samples are required.

- The field personnel will prepare a "Sample Request Form" indicating the need for formal chain of custody procedures.

4.2  Sample Containers

4.2.1  Container Material  (See also §7.2)

- Samples that are to be sent to an outside laboratory for analysis should be collected in appropriate containers supplied by that laboratory.

- Most samples for in-house analysis can be collected in either glass (borosilicate) or plastic (polyethylene, conventional or linear polypropylene).

4.2.2 Container Preparation

4.2.2.1  General procedure for all containers

- Inspect containers.  Discard if deteriorated or if contaminated to the extent that they cannot be cleaned.

- Rinse well with tap water.

- Wash containers and caps using a non-phosphate detergent, warm water and a brush.  Alternatively, the dishwasher may be used if appropriate.

- Rinse well with tap water.

- Rinse several times with portions of deionized water.

- Invert and let dry.  Re-inspect before storage.

4.2.2.2  Containers for Metals Analyses

- Clean using the General procedure above.

- Soak container with dilute nitric or hydrochloric acid at least overnight.  (May be stored longer.)

- Rinse at least five times with deionized water. Invert and air dry.

- Inspect and label with date and I.D. before storage.

      - Containers for Bacteriological Samples

- Clean using the General procedure above.

- Add sodium thiosulfate and sterilize before storage.

- Commercially available sterilized and preserved may be used.

4.2.2.4  Containers for Organic Samples

- Clean glass containers using the procedure for metals samples.

- Rinse at least twice with Pesticide quality acetone, methanol, or hexane.

- Air dry.  Residual vapors should be removed by using vacuum to draw air through the container.

4.3  Sample Preservation and Storage  (See also §7.2)

4.3.1  Required monitoring samples

Use the information published as "Table II - Required Containers, Preservation Techniques, and Holding Times" at 40 CFR 136.3

4.3.2  Other samples

See reference cited in 4.3.1.  Alternatively, use Standard Methods, 18th Edition, Part 1060 C., APHA-AWWA-WPCF, 1992.

5.  Sample Custody

Statement of Policy

It is the policy of the OMI/TW Laboratory that a sample collected by an employee normally assigned to work in the Laboratory is considered to be in the custody of the Laboratory at the moment of collection.  Delivery of the sample to the laboratory does not constitute a change of custodian.  Receipt of the sample in the Laboratory is documented on the Sample Log; a Chain of Custody Form is not required.

The transfer of the custodianship of a sample between the Laboratory or an employee of the Laboratory, and a non-Laboratory employee or entity shall be documented.  The method of documentation shall be appropriate for the type of transfer.  Whatever method that is used shall contain as a minimum clear identification of the sample, and the dated and timed signatures of the relinquisher and receiver.

5.1  Sample Custody - A sample is under custody if:

- It is in your possession, or

- It is in your view, after being in your possession, or

- It was in your possession and then was placed in a secured location to prevent tampering, or

- It is in a designated secured area.

  - Routine Samples

Most samples handled by the Laboratory are routine in nature.  Such samples are checked for integrity and proper labeling, and are logged onto the daily sample log.  The sample log is a permanent record that includes the sample source; the dates, times, and identification of the personnel associated with sample collection and receipt; and other information characterizing the sample and methods of preservation.

5.3  Special Samples - Non-routine/special samples are either collected by laboratory personnel or  will be accompanied by a chain of custody form which will include pertinent information describing the sample and the required analyses.  This form may be provided either in-house for work to be done in-house, or by an outside laboratory if the analyses are to be contracted out and the sample merely passing through the laboratory.

- 5.3.1  In-house Chain of Custody

- A form entitled “Custody Transfer Record/Laboratory Analysis Request” (or an alternate form) is provided by the sampler, identifying the sample(s) and documenting collection time and the testing required.

- The sample receiver inspects, receives, and logs the sample(s) after insuring that all required information has been provided.

- 5.3.2  Transfer of Custody

- When the possession of the sample is transferred, the individuals relinquishing and receiving will sign, date, and note the time on the record.

- If the sample is to be shipped to another lab it will be accompanied by a separate custody record provide by the other laboratory and will be packaged and shipped in a manner appropriate for the analyses required.

- 6.  Calibration Procedures and Frequency

6.1 Meters - Dissolved oxygen, electrical conductivity, and pH meters are standardized before use and their performance verified after measurements are made.  The values, time and operator's I.D. are recorded on worksheets associated with the tests.  Dissolved oxygen meters should be checked by the Winkler dissolved oxygen method weekly.

6.2  Temperature measurement devices are calibrated at least annually against a reference thermometer.  Mercury thermometers will be used in the laboratory.  Spirit thermometers or electronic temperature measuring devices may be used in the field.

6.3  Personnel will be allowed to use the laboratory equipment and instruments only after they have been trained.

7.  Analytical Procedures

7.1  Included References

Included in this document by reference are the following publications:

Handbook for Analytical Quality Control in Water and Wastewater Laboratories, EPA-600/4-79-019, 1979

- Methods for Analysis of Water and Wastes, EPA-600/4-79-020, 1979 (Revised 1983)

Methods for the Determination of Inorganic Substances in Environmental Samples, EPA/600/R-93/100, August 1993

Methods for the Determination of Metals in Environmental Samples, EPA/600/4-91/101, June 1991

Standard Methods for the Analysis of Water and Wastewater, 18th Edition, APHA-AWWA-WPCF, 1989

- Microbiological Methods for Monitoring the Environment, EPA-600/8-78-017, 1978

- Methods for E. coli in Drinking Water, EPA/600/4-91/016, July 1991

Methods for Measuring the Acute Toxicity of Effluents to Freshwater and Marine Organisms (Third Edition), EPA-600/4-85/013 (Third Edition)

## Methods for Measuring the Acute Toxicity of Effluents to Freshwater and Marine Organisms (Third Edition), EPA 600/4-90/027F (Fourth Edition)

Only those methods which have been approved by a regulatory agency and for which the laboratory is accredited will be used for regulatory compliance.  Any method may be used for solely internal, non-regulatory purposes (e.g.- process control).

  - Methods, Containers, Preservation, and Holding Times

The following table summarizes analytical methods used in the Laboratory, and includes requirements for appropriate containers, methods for sample preservation, and holding times.

Methods – Preservation – Hold Times

| FoT | Analyte / Method Description | Method # |  | Container |  | Preservation |  | Hold Time |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FoT 101 MICROBIOLOGY OF DRINKING WATER |  |  |  |  |  |  |  |  |  |
| Microbiology of Drinking Water |  |  |  |  |  |  |  |  |  |
| 101.010 | Heterotrophic Plate Count | SM 9215 B |  | P, G |  | Cool, 4C, 0.008% Na2S2O3 |  | 30 / 6 hours |  |
| 101.060 | Total Coliform by Colilert (Presence/Absence) | SM 9223 |  | P |  | Cool, 4C, 0.008% Na2S2O3 |  | 30 hours |  |
| 101.060 | E. coli by Colilert (Presence/Absence) | SM 9223 |  | P |  | Cool, 4C, 0.008% Na2S2O3 |  | 30 hours |  |
| Microbiology of Drinking Water (Enumeration/Source) |  |  |  |  |  |  |  |  |  |
| 101.120 | Total Coliform by Multiple Tube Fermentation (MTF) | SM 9221 A,B,C |  | P, G | Cool, 4C, 0.008% Na2S2O3 |  | 8 hours |  |  |
| 101.130 | Fecal Coliform by EC Media from MTF | SM 9221 E |  | P, G | Cool, 4C, 0.008% Na2S2O3 |  | 8 hours |  |  |
| FoT 107 MICROBIOLOGY OF WASTEWATER |  |  |  |  |  |  |  |  |  |
| 107.020 | Total Coliform by Multiple Tube Fermentation (MTF) | SM 9221 B |  | P, G | Cool, 4C, 0.008% Na2S2O3 |  | 6 hours |  |  |
| 107.040 | Fecal Coliform by EC Media from MTF | SM 9221 C, E (MTF/EC) |  | P, G | Cool, 4C, 0.008% Na2S2O3 |  | 6 hours |  |  |
| FoT 102 INORGANIC CHEMISTRY/PHYSICAL PROPERTIES OF DRINKING WATER |  |  |  |  |  |  |  |  |  |
| 102.100 | Alkalinity – titration | SM 2320 B |  | P, G |  | Cool, 4C |  | 14 days |  |
| 102.170 | Chloride – titration, argentometric | SM 4500-Cl- B |  | P, G |  | none required |  | 28 days |  |
| 102.163 | Chlorine, Total – DPD colorimetric (Hach) | SM4500-Cl G |  | P , G |  | Analyze immediately |  | none |  |
| 102.200 | Fluoride – ISE | SM 4500-F C |  | P |  | none required |  | 28 days |  |
| 102.121 | Hardness – titration | SM 2340 C |  | P, G |  | HNO3 to pH < 2 |  | 6 months |  |
| 102.140 | Total Filterable Residue (TDS) – gravimetric | SM 2540 C |  | P, G |  | Cool, 4C |  | 7 days |  |
| 102.130 | Conductivity – specific conductance | SM 2510 B |  | P, G |  | Cool, 4C |  | 28 days |  |
| (Note: SM refers to the 18th Edition of Standard Methods for the Analysis of Water and Wastes) |  |  |  |  |  |  |  |  |  |
| FoT | Analyte / Method Description | Method # |  | Container |  | Preservation |  | Hold Time |  |
| FoT 108 INORGANIC CHEMISTRY OF WASTEWATER |  |  |  |  |  |  |  |  |  |
| 108.410 | Alkalinity – titration | SM 2320 B |  | P, G |  | Cool, 4C |  | 14 days |  |
| 108.495 | Ammonia – distillation, titration | SM 4500-NH3 B,E |  | P, G |  | Cool, 4C, H2SO4 to pH < 2 |  | 28 days |  |
| 108.590 | BOD – electrode | SM 5210 B |  | P, G |  | Cool, max. 4C |  | 24 hours |  |
| 108.591 | CBOD – electrode w/ nitrification inhibitor | SM 5210 B |  | P, G |  | Cool, max. 4C |  | 24 hours |  |
| 108.450 | Chloride – titration, argentometric | SM 4500-Cl- B |  | P, G |  | Cool, 4C |  | 28 days |  |
| 108.464 | Chlorine Residual, Total – titration, DPD-FAS | SM 4500-Cl F |  | P, G |  | Analyze immediately |  | none |  |
| 108.421 | Hardness – titration, EDTA | SM 2340 C |  | P, G |  | HNO3 to pH < 2 |  | 6 months |  |
| 108.495 | Kjeldahl Nitrogen – digestion, distillation, titration | SM 4500-Norg B SM 4500-NH3 B,E |  | P, G |  | Cool, 4C, H2SO4 to pH < 2 |  | 28 days |  |
| 108.531 | Oxygen, dissolved – electrode | SM 4500-O G |  | P, G |  | Analyze immediately |  | none |  |
| 108.490 | pH – electrometric | SM 4500-H+ B |  | P, G |  | Analyze immediately |  | 0.5 hour |  |
| 108.440 | Residue, Total – gravimetric | SM 2540 B |  | P, G |  | Cool, 4C |  | 7 days |  |
| 108.441 | Residue, Filterable (TDS) – gravimetric | SM 2540 C |  | P, G |  | Cool, 4C |  | 7 days |  |
| 108.070 | Residue, Nonfilterable (TSS) – gravimetric | EPA 160.2 |  | P, G |  | Cool, 4C |  | 7 days |  |
| 108.443 | Residue, Settleable – volumetric | SM 2540 F |  | P, G |  | Cool, 4C |  | 7 days |  |
| 108.090 | Residue, Volatile – gravimetric | EPA 160.4 |  | P, G |  | Cool, 4C |  | 7 days |  |
| 108.430 | Specific Conductance | SM 2510 B |  | P, G |  | Cool, 4C |  | 28 days |  |
| 108.390 | Turbidity – nephelometric | SM 2130 B |  | P, G |  | Cool, 4C |  | 48 hours |  |
| FoT 118 Whole Effluent Toxicity of Wastewater |  |  |  |  |  |  |  |  |  |
| 113.010 | Fathead Minnow, Continuous Flow, Survival in 100% | EPA 600/4-90/027F |  | N/A |  | N/A |  | N/A |  |
| 113.021 | Fathead Minnow, Continuous Flow, Survival in 100% | EPA 2000 (EPA-821-R-02-012) |  | N/A |  | N/A |  | N/A |  |
| (Note: SM refers to the 18th Edition of Standard Methods for the Analysis of Water and Wastes) |  |  |  |  |  |  |  |  |  |

- 8.  Data Reduction, Validation, and Reporting

8.1  Data Reduction - Original data and calculated results are recorded by the analyst in ink on pre-printed worksheets.  Calculations for each method are established in the procedures.  Where possible, calculations are performed by entering data into a formatted spreadsheet, where the calculations are made by spreadsheet formulas and macros.  After calculations are made, the worksheet is reviewed and signed by the analyst and is submitted for further review.

8.2  Data Validation

8.2.1  Quality assurance review - The lead person of each section thoroughly reviews each worksheet for completeness, accuracy, and compliance with quality control requirements.  The final result for each analysis is determined.  Unacceptable data is rejected.  Samples with unacceptable data will be reanalyzed whenever possible.  The lead person signs the approved worksheet and submits it for final review/approval by the Laboratory Supervisor.

- 8.2.2  Final Approval - The Laboratory Supervisor reviews all worksheets and signs them when they are approved.

8.3 Reporting of results

Unlike a commercial laboratory where data is reported to outside clients, all data generated by the laboratory is reported internally.  For this reason, worksheets signed by the Laboratory Supervisor are considered final reports.  Each report (worksheet) contains the original data, calculated results, quality control sample results, and signatures indicating acceptance and approval.

Data generated by the laboratory is generally reported in summary format, either by submission of monthly reports to regulatory agencies, or entry into a database.

9.  Internal Quality Control Checks

9.1  Types and purposes of method QC procedures

9.1.1  Method (Reagent) Blank - provides a check of interferences which may be introduced during the analysis.  The sample usually consists of laboratory pure water carried through all steps of the analytical procedure.

9.1.2  Duplicate (Replicate) Sample -  one or more separate portions of a sample analyzed in parallel.  Comparison of data from all replicates will provide a measure of the precision of the analysis under a specific set of conditions.

When performed by several laboratory personnel, the data will provide a measure of the precision of the method to be used for determining the control limits for evaluating the analytical results.

9.1.3  Method Spike - a method blank to which a known concentration of the analyte(s) has been added.  This sample is carried through all steps of the analytical procedure.  Use of this sample will provide a measure of the accuracy of the analysis (percent recovery) and consequently a measure of the validity of the analytical results.

9.1.4  Sample Spike - a sample to which a known concentration of the analyte(s) has been added.  The concentration of the added analyte should be from one to five times the level in the sample or five to ten times the analyte detection limit.  The added analyte should not cause the total concentration to exceed the working range of the method so as to require a dilution of the fortified sample.

Analysis of a sample spike will provide a measure of sample matrix interferences for a specific type of sample.  Each type of sample must be evaluated individually and must be reevaluated when the character of the sample changes.

9.1.5  Reference Sample - a sample purchased or otherwise obtained from an outside agency or vendor having has a known analyte concentration.  It is analyzed as a regular sample for comparison with its known value.  Since the analyst does not know the expected value, this will provide an unbiased evaluation of the test procedure.  Reference samples may be used to prepare fortified samples.

9.1.6  Check Sample - a sample of known analyte concentration purchased or prepared in-house to be run on a routine, day-to-day basis.

9.1.7  Reporting Limit Verification Sample (RLV) -  a sample of known analyte concentration run as needed to verify analyte recovery at the predetermined reporting limit.

9.2  Types and purposes of sample QC procedures

9.2.1  Container Blank - a sample of laboratory pure water which has been held in a prepared sample container and does not leave the laboratory.  Analysis of this sample will provide a measure of the efficiency of the container preparation procedure and technique.

9.2.2  Trip Blank - a sample of laboratory pure water placed in a sample container that is transported with and handled in the same manner as other containers while samples are being collected and returned to the laboratory.  Analysis of the trip blank will provide a measure of the extent to which samples may become contaminated in the field and in the laboratory before analysis.

9.2.3  Field Blank - a sample of laboratory pure water which has been carried through sample collection devices after decontamination.  To the extent that it is possible, all sample collection procedures should be evaluated.  Analysis of this sample will provide a measure of the extent to which a sample may be contaminated by the sampling equipment.

9.2.4  Field Duplicate - a separate (second) sample collected at the sampling site.  It will provide a measure of the homogeneity of the sample material and the consistency of sampling technique.

9.3  Analytical Reagents

9.3.1  Chemicals used for analyses are purchased from various vendors.  All are of "Reagent Grade" or better.  Solvents used for preparation of organic sample containers are pesticide grade.  Acids for metals analyses will be "For Trace Metal Analysis” or better.

9.3.1  Purified water for general laboratory use (e.g. - reagent preparation, final rinse of glassware washing procedures) is tap water that has been passed through a reverse osmosis membrane, followed by mixed-bed deionization canisters, UV sterilization and polishing filters.  It is continuously circulated in a loop throughout the laboratory and is available distilled in an all-glass still as laboratory pure water.

9.3.2  Distilled water - Purified water is distilled in an all-glass still for use preparing bacteriological media, and for preparing standards and samples for trace metals analyses.

9.4  Standards

9.4.1  Metal and fluoride standards are purchased prepared and are replaced before the expiration date printed on the label, or within one year from receipt if no expiration date.

9.4.2  pH Buffers are purchased prepared and are traceable to NBS, and are replaced before the expiration date printed on the label.

9.4.3  Other standards are prepared as needed from primary standards or appropriate quality reagents.

9.5  Methods - Approved/accredited (regulatory) or appropriate (process control) methods are chosen.  Each method has appropriate quality control checks with acceptance limits.  These checks may be contained within the original method itself, or may be determined within the laboratory.

9.6  Recording of data

- All analytical data are recorded in bound notebooks or on preprinted worksheets and become a permanent record in the laboratory.

- All printed data generated by an instrument will be dated, annotated so as to clearly identify samples, and will be kept as original data.

- The analyst's I.D. will accompany all data.

- All original data are retained for at least five years.

9.7  Calculation and reporting of analytical results

- Results of analytical determinations will be made by equations contained within the written established methods that are used.

- Results will be rounded to the correct number of significant figures after all calculations have been made.

- The correct number of significant figures will be determined by evaluation of all sources of error and the precision of each analysis.

9.8  Review of analytical results

All analytical results will be reviewed/approved by the Chemist/Microbiologist to verify computations accuracy of calculations and to see that the data meet the established criteria.  Final approval of the data will be by the Laboratory Supervisor.

9.9 Acceptable limits of analytical results

- In general, analysis results are rejected if duplicates are not within 10 %, and standards and spikes are not within range.

- The acceptable range for standards and spikes is determined by evaluation of recent historical data for each analysis.

- Specific controls and acceptance criteria may be established for each individual method. (See table next page.)

- If possible, samples are reanalyzed after remedial action is taken.

| Quality Control Acceptance Criteria |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FoT | Analyte | Method |  | %S 1 | %R 2 | Method Blank | Other |  |
| 102.100 | Alkalinity - titration | SM 2320 B |  | na | 10% | na |  |  |
| 102.170 | Chloride - titration, argentometric | SM 4500-Cl- B |  | 80-120% | 10% | na |  |  |
| 102.163 | Chlorine, Total - DPD colorimetric (Hach) | SM4500-Cl G |  | na | na | na |  |  |
| 102.200 | Fluoride - ISE | SM 4500-F C |  | 80-120% | <1: 25%; >1: 10% | na | CV/QCS: 90-110% |  |
| 102.121 | Hardness - titration | SM 2340 C |  | 80-120% | 10% | na | LCS: 80-120% |  |
| 102.140 | Total Filterable Residue (TDS) - gravimetric | SM 2540 C |  | 90-110% | na | na |  |  |
| 102.130 | Specific conductance | SM 2510 B |  | 980-1020 µS | na | na |  |  |
| 108.410 | Alkalinity - titration | SM 2320 B |  | na | 10% | na |  |  |
| 108.495 | Ammonia - distillation, titration | SM 4500-NH3 B,E |  | 80-120% | <2: 25%; >2: 10% | na |  |  |
| 108.590 | BOD - electrode | SM 5210 B |  | GGA: 198±30 mg/L | na | Depletion < 0.2 mg/L | DOf ≥ 1.0 mg/L; Min. Depl.=2 mg/L |  |
| 108.591 | cBOD - electrode w/ nitrification inhibitor | SM 5210 B |  | na | na | na | DOf ≥ 1.0 mg/L; Min. Depl.=2 mg/L |  |
| 108.450 | Chloride - titration, argentometric | SM 4500-Cl- B |  | 80-120% | 10% | na |  |  |
| 108.464 | Chlorine Residual, Total - titration, DPD-FAS | SM 4500-Cl F |  | na | na | na |  |  |
| 108.421 | Hardness - titration, EDTA | SM 2340 C |  | 80-120% | 10% | na | LCS: 80-120% |  |
| 108.495 | Kjeldahl Nitrogen - digestion, distillation, titration | SM 4500-Norg B SM 4500-NH3 B,E |  | 80-120% | <2: 25%; >2: 10% | na |  |  |
| 108.531 | Oxygen, dissolved - electrode | SM 4500-O G |  | na | na | na | Calibration ± 0.2 mg/L |  |
| 108.490 | pH - electrometric | SM 4500-H+ B |  | na | na | na | Standard ± 0.05 s.u. |  |
| 108.440 | Residue, Total (TS) - gravimetric | SM 2540 B |  | na | na | na | % Range = 10*(n-1); min. residue 1 mg |  |
| 108.441 | Residue, Filterable (TDS) - gravimetric | SM 2540 C |  | 90-110% | na | < 5 mg/L | % Range = 10*(n-1); min. residue 1 mg |  |
| 108.070 | Residue, Nonfilterable (TSS) - gravimetric | EPA 160.2 |  | 90-110% | na | na | % Range = 10*(n-1); min. residue 1 mg |  |
| 108.443 | Residue, Settleable (SS) - volumetric | SM 2540 F |  | na | na | na |  |  |
| 108.090 | Residue, Volatile (VS) - gravimetric | EPA 160.4 |  | na | na | na | % Range = 10*(n-1); min. residue 1 mg |  |
| 108.430 | Specific Conductance (EC) | SM 2510 B |  | 980-1020 µS | na | na |  |  |
| 108.390 | Turbidity - nephelometric | SM 2130 B |  | na | na | < 0.05 NTU | 2° Standards: ±10% of assigned value |  |
|  | 1 - Percent recovery of a standard (blank spike, sample spike, control sample) |  |  |  |  |  |  |  |
|  | 2 - Relative percent difference between replicates (sample duplicate, MS/MSD) |  |  |  |  |  |  |  |
|  | Revised 1/30/2006 |  |  |  |  |  |  |  |

9.10  Definitions of analytical limits

9.10.1  Method Detection Limit (MDL) - The minimum quantity that can be reported with a 99% confidence limit that the analyte concentration is above zero.  This is taken to be the appropriate Student’s t value times the standard deviation of a minimum of 7 replicate samples at the detection limit (from 40 CFR 136, Appendix B).

9.10.2  Instrument Detection Limit (IDL) - The minimum signal detectable above the background noise level.  This usually is determined as 2-3 times the standard deviation at the detection limit.

    - Reporting Limit (RL) - The lowest analyte concentration which can be reported with demonstrated adequate recovery.

    - Minimum Level (ML) - The analyte concentration in the sample which corresponds to the lowest valid standard concentration, taking into account dilution/concentration of the sample during the analytical procedure. (functional definition)

9.10.5  Spike Recovery (%S) - The ratio of the measured amount of an analyte added to the know amount added, multiplied by 100%.

9.10.6  Percent Range (%R) - The difference between the highest and lowest values, divided by the mean of the measurements, multiplied by 100%.

9.11  General Requirements

9.11.1  Reagent Blank is analyzed with each set of samples where appropriate.

9.11.2 Standards are run with each set of samples as appropriate to the method.

9.11.3  Sample controls

(1) Ten percent of each type of sample in a batch (minimum one per batch) should be run in duplicate.  All solids samples (e.g.- total suspended solids, total solids, etc.) are run in at least duplicate.

(2)  Blank spike if practical.

(3)  Five to ten percent of samples should be spiked if practical.

(4)  One or more check samples if available.

(5)  Reporting Limit Verification sample if appropriate.

9.11.4  Performance/maintenance records are kept for the various instruments.

9.11.5  Data is recorded, and all charts and print-outs are retained in some format as a permanent record.

9.11.6  Calculations are documented and checked by QA/QC.

9.11.7  Non-routine samples are monitored individually throughout the analytical procedure and are evaluated for matrix interferences.  Any unusual occurrences are brought to the attention of the Chemist and Laboratory Supervisor.  If a procedure must be modified or an alternate procedure used, this fact will be noted in the analysis records and on the report of the analytical results.

9.12  Bacteriology

- Positive and negative controls on media.

- Water suitability test (annually).

- Inhibitory residue test (as needed).

- Sterility checks on media, agar, sample bottles and dilution water.

- Trace metals in water used for media preparation and sample dilution.

- Recording thermographs are used to monitor the performance of incubators.

9.13  Biotoxicology

- Samples are run in duplicate.

- Duplicate controls are run with each sample.

- A reference toxicant is run with each batch of fish.

- 10.  Performance and System Audits

The Laboratory Supervisor will on a regular basis review the performance of the laboratory to insure that the Quality Assurance program contains all essential elements, and that all requirements of the Quality Assurance program are being met.  Any inadequacies will be evaluated and corrected immediately by whatever means are appropriate.

This document will from time to time be reviewed to insure that the Quality Assurance program is relevant to the work done in the laboratory, that it is sufficient to insure the quality of the data reported by the laboratory, and that it meets the requirements of the accrediting  agency.

From time to time, various types of audits are performed by corporate quality assurance personnel and deficiencies are reported to the laboratory.  Audits are made of record keeping, data integrity, reporting accuracy, and documentation.

- 11.  Preventive Maintenance

The performance of all equipment and instruments will be monitored continuously.  Any indication of improper operation will be reported to the Laboratory Supervisor for remedial action.  All major adjustments or repairs will be performed by manufacturer’s authorized service personnel.

11.1  General laboratory equipment is maintained on an as needed basis, but at least semiannually.  A maintenance log will be prepared for each piece of equipment which will contain:

- The equipment operation manual,

- A summary of required and recommended maintenance tasks,

- Documentation of when, and by whom, maintenance tasks were performed,

- Copies of field service reports for maintenance activities conducted by an outside agency (e.g.- by a commercial repair service or by a manufacturer's field service representative under an annual maintenance contract), and

- Documentation of any remedial actions.

11.2  Maintenance Contracts - The analytical balances, water purification system, and the autoclave are on annual maintenance contracts for manufacturer supplied maintenance.

12.  Assessment of Precision and Accuracy

12.1  Precision is a measurement of the reproducibility of a measurement, and may be measured by the standard deviation or relative range of a group of measurements, or the relative percent difference between two measurements.

12.1.1 Standard deviation is a statistical calculation which may be used, for a normally distributed set of measurements, to estimate the confidence interval around a measured value.  A major use of this parameter is in determining acceptance limits for the measurement of a specific parameter over time (control charts)

12.1.2 Relative range, or difference, is a simple calculation used for a limited set of data to determine variability of results.  Limits are determined generally from historical data, and are dependent on the specific method and sample type.

12.2  Accuracy is measured by analyzing materials of known composition and determining the  efficiency of recovery.

12.2.1  Checks are made by spiking samples and blanks, and by running check samples.

12.2.2  Reference samples should be run at least four times per year.  These are samples from outside the laboratory which have known compositions, or for which the analyses are evaluated (e.g.- EPA performance samples).

- 13.  Corrective Actions

The Laboratory Supervisor will be notified when a problem occurs with the analysis of control samples.  Under the direction of the Laboratory Supervisor the following steps will be taken:

- Determine if the appropriate analytical method was used.

- Verify that the standards and reagents used have not expired and/or are of suitable quality.

- Verify that the analyst used the correct analytical techniques.  Retrain analyst in proper techniques

- Determine if the equipment or instrument was functioning correctly.  Have replaced or repaired as necessary.

- Have the analyst re-run the procedure.

- Have a second analyst run the procedure.

- Re-evaluate the implementation of the procedure in the laboratory.

- 14.  Quality Assurance Reports

Quality assurance reports are not generally published.  Specific sample data reporting is internal.  Worksheets with final approval are considered final reports of individual data.  Each worksheet contains the results of applying the quality assurance process.

No reports are generated relative to the review or auditing of the quality assurance program.  Inadequacies and corrections are addressed as needed.