THE SAN DIEGO STREAM TEAM
QUALITY ASSURANCE PROJECT PLAN

COMPLETED PLAN PREPARED BY:

Lisa Homna

Isabelle Kay

David Gibson

Andre Macedo

Cynthia Mallett

Barry Lindgren

Juliette Nowak

__________________________

Refer correspondence to:

Barry Lindgren

barry@canyontrace.com

(760) 436-7991

Approvals:

Agency / Organization:  ________________________________________________

Signature:  __________________________________________Date: ____________

Agency / Organization:  ________________________________________________

Signature:  __________________________________________Date: ____________

Agency / Organization:  ________________________________________________

Signature:  __________________________________________Date: ____________

# 2.	Table of Contents   								page

1. Title and Approval Page	1

2. Table of Contents	2

3. Distribution List	4

4. Project Organization	4

5. Problem Definition/Background	4

6. Project/Task Description	6

7. Data Quality Objectives	7

8. Training Requirements	10

9. Documentation and Records	10

10. Sampling Process Design	11

11. Sampling Method Requirements	12

12. Sample Handling and Custody Procedures	12

13. Analytical Methods Requirements	13

14. Quality Control Requirements	14

15. Instrument/Equipment Testing, Inspection and Maintenance	16

16. INSTRUMENT Calibration / Standardization and frequency	17

17. Inspection/Acceptance Requirements	18

18. Data Acquisition Requirements	18

19. Data Management	18

20. Assessment and Response Actions	18

21. Reports	19

22. Data Review, Validation and Verification	19

23. Validation and Verification Methods	19

24. Reconciliation with DQOs	19

Appendices

Data Quality Forms 								Appendix 1

Data and Observation Sheets 							Appendix 2

Maps of Sampling Sites								Appendix 3

# 3.	Distribution List

All  field team leaders, equipment managers, data managers, quality assurance leaders, and technical advisors will receive copies of this Quality Assurance Project Plan (QAPP), and any approved revisions of this plan.  Once approved, this QAPP will be available to any interested party by requesting a copy from Barry Lindgren.

# 4.	Project Organization

The purpose of this organization will be to assess the physical, chemical and biological condition of San Diego County rivers and streams, educate students and citizens on biological monitoring, stream ecology, and restoration, and contribute data to the statewide effort to assess the health of California rivers and streams. SDST will monitor and assess inland surface waters in the county annually:

Santa Margarita River watershed

San Luis Rey River watershed

Carlsbad Hydrologic Unit watersheds

San Dieguito River watershed

Los Penasquitos Hydrologic Unit watersheds

San Diego River watershed

Sweetwater River watershed

Otay River watershed

Tijuana River watershed

and biannually in:

San Juan Creek watershed

San Felipe Creek watershed

Anza Borrego watershed

4.1 Management (Director, Field Coordinators, Laboratory Coordinators)

- Barry Lindgren

- Cynthia Mallet

- Lisa Homna

- Lan Wiborg

4.2 Field Team Leaders

- Andre Macedo

- Cynthia Mallett

- Deborah LeLevier

- Isabelle Kay

- Lan Wiborg

4.3 Data and Equipment Managers

- Lilian Busse

- Lisa Homna

- Rob Roy

4.4 Quality Assurance Personnel

- Lan Wiborg

- Lisa Homna

- Barry Lindgren

- Jayne Strommer

- Isabelle Kay

- Dr. Neal Biggart

4.5 Technical Advisors

- David Gibson, Environmental Specialist III, Regional Water Quality Control Board

- Andre Macedo, Entomologist/Vector Ecologist, City of San Diego Environmental Monitoring and Technical Services Division

- Isabelle Kay, Natural Reserve System Manager, University of California, San Diego

# 5.	Problem Definition/Background

## 5.1.	Problem Statement

There is insufficient information to adequately assess the status of aquatic resources in San Diego Region watersheds.  There are concerns over the status and maintenance of the quality of water resources in these watersheds.  Citizen monitoring organizations have been formed in local watersheds to address these water quality concerns.  If quality assurance is adequate, valuable information will be provided for watershed management and pollution prevention.

#### 5.1.1.1.	Mission

The San Diego Stream Team (SDST) is a citizen monitoring team dedicated to the bioassessment of San Diego County rivers and streams using State and Federal protocols and to educating the public about the biological condition of the local waters. The SDST meets at least quarterly and conducts bioassessment sampling in Spring and Fall to assess the condition of San Diego rivers and streams.

Proper sampling and analysis of benthic communities can result in rapid, inexpensive, easily communicated, and accurate assessments of stream health. Quarterly seminars and up to two training workshops a year help members stay current with developments in biological monitoring.

Membership is open to anyone interested in stream ecology and biology. Trial members, teachers, students, natural resource management professionals, and private citizens can all benefit from SDST membership and activities. Members receive training and support for bioassessment projects and access to equipment and supplies.

The SDST also participates in educational events, stream clean-ups, and habitat restoration. The SDST is actively seeking new members and partnerships with San Diego Tribes, agencies, and other citizen groups.

#### 5.1.1.2.	Program Goals

The general goals of citizen monitoring are:

- Identifying the status and trends of biological resources in and around the local aquatic environment,

- Screening for water quality problems,

- Identifying pollution sources and illegal activities (spills, wetland fill, diversions, discharges),

- Establishing trends in water quality for waters that would otherwise not be monitored,

- Evaluating the effectiveness of restoration or management practices,

- Evaluating the effect of a particular activity or structure, and

- Evaluating the quality of water compared to specific water quality criteria,

- Developing a body of knowledge about the unique character of San Diego streams with respect to insect and stream ecology.

In addition, citizen monitors build awareness of water quality issues, aquatic resources and pollution prevention.

This project will supplement existing agency information by monitoring streams in San Diego Region watersheds.  The focus of the project is on habitat and chemical, physical and biological water quality measures that will identify the status of these aquatic resources.  The information obtained will posted on the San Diego Stream Team website and wil also be provided to the regulatory agencies.  It is the responsibility of the regulatory agencies to ensure that adequate and valid data are collected to meet their regulatory requirements.

## Intended Usage of Data

The data will be used by the San Diego Stream Team for general watershed assessment purposes.  This assessment of this data will be useful in providing information for watershed management and pollution prevention.  The data will be made available to the public for purposes of watershed education.  It will also be made available to the regulatory and resource management agencies to supplement their existing data collection efforts.  One potential application of the data will be to provide information to the Regional and State Boards for their use in Section 305(b), 303d, and TMDL reporting.

Data will be compiled and maintained at the San Diego Stream Team laboratory. The information will be shared with the State Water Resources Control Board, the San Diego Regional Water Quality Control Board, and upon request to other state, federal, and local agencies and organizations.

# 6.	Project/Task Description

## General Overview of Monitoring

The citizen monitors are monitoring water quality in the San Diego Regional Water Quality Control Board watersheds.  Table 6.1 summarizes the monitoring design, including the physical, chemical and biological parameters to be measured and whether the samples will be analyzed by the monitoring group in the field or retained for later analysis in the Stream Team laboratory or by a professional laboratory.

Table 6.1   Summary of Field Monitoring Design

| PARAMETER | TYPE OF MONITORING |
| --- | --- |
| Flow | F |
| Temperature | F |
| Dissolved Oxygen | F |
| pH | F |
| Conductivity (fresh water) or Salinity (marine) | F |
| Nitrate | F |
| Benthic Macroinvertebrates | L, P |
| Odor and Visual Observations | F |
| Phosphate | F |

Codes for Table 6.1

Type: F:  field analysis,   L: in-house lab analysis,   P:  sample only, send to outside professional lab

All of the water quality data will be compared to the Regional Water Quality Control Board Basin Plan.  For results that are not comparable to the Basin Plan we will review those data with our Technical Advisors.

This QA plan only addresses data quality objectives for the following parameters:

- Flow

- Temperature

- Dissolved Oxygen

- pH

- Conductivity

- Phosphate (phosphorus)

- Nitrate (nitrogen)

- Benthic Macroinvertebrates

For stream and urban storm drain environments, flow will be determined by using the protocol described in the Measuring the Health of California Streams and Rivers Manual.  Chemistry and physical parameters will be monitored using protocols outlined in the Measuring the Health of California Streams and Rivers Manual.   Benthic macro-invertebrate monitoring will be performed according to the California Stream Bioassessment Procedure.  This program has a systematic method for visual and other sensory observations. Observational data include color, odor, presence of oil or tar, trash, foam, and algae.  In addition, the stream habitat quality will be assessed, at least twice per year, depending on flow, using the California Dept. of Fish and Game Physical Habitat Assessment Form.  Observational data include epifaunal substrate/available cover, embeddedness, velocity/depth regimes, sediment deposition, channel flow status, channel alteration, frequency of riffles, bank stability, vegetative protection, and riparian vegetative zone width.

## .

Table 6.2 identifies the schedule of major activities associated with this project.

Table 6.2  Project Schedule

| Activity | Date |
| --- | --- |
| Identify monitoring leaders | March and September annually |
| Obtain training for monitoring leaders | March and September annually |
| Recruit monitors | March and September annually |
| Obtain and check operation of instruments | March and September annually |
| Train monitors | March and September annually |
| Initiate monitoring | May and October annually |
| Initiate data entry | January and July annually |
| Calibration and quality control sessions | April and August annually |
| Review data with technical advisors | April and August annually |

# 7.	Data Quality Objectives

This section identifies how accurate, precise, complete, comparable, sensitive and representative measurements will be.  These data quality objectives were derived by reviewing the QA plans and performance of other citizen monitoring organizations (e.g. Chesapeake Bay, Texas Watch, Coyote Creek Riparian Station, Southern California Citizen Monitoring Steering Committee, Heal the Bay Malibu StreamTeam), by considering the specifications of the instruments and methods which we will employ, and by considering the utility of the data.  For purposes of this QAPP the data quality is considered adequate for the determination of general water quality conditions, with a potential application of the data to Section 305(b) and 303(d) reporting purposes.

Data quality objectives are summarized in Tables 7-1 to 7-3.  Whenever possible the methods with the greatest sensitivity and lowest detection limit will be employed as the primary methods.  Methods with lesser sensitivity and higher detection limits will be used for field confirmations or as back-up methods in the case that the primary methods are not available or functioning properly for a particular sampling event.

Table 7.1.  Data Quality Objectives for Conventional Water Quality Parameters

| Parameter | Method/range | Units | Detection Limit | Sensitivity* | Precision | Accuracy | Complete-ness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature | Thermometer (-5 to 50) | o C | -5 | 0.5 o C | ± 0.5 o C | ± 0.5 o C | 80% |
| Dissolved oxygen | Vacuum ampoule Indigo carmine | mg/l | 1.0 mg/l | 1.0 (1.0-6.0) 2.0 (6.0-12.0) | + 1.0 mg/l | + 1.0 mg/l | 80% |
| pH | pH meter | pH units | 2.0 | 0.1 unit | ± 10% | ± 10% | 80% |
| pH | Non-bleeding strips (range 4.5-10.0) | pH units | 4.5 | 0.5 unit | + 0.5 units | + 0.5 units | 80% |
| Conductivity | conductivity meter | µS/cm | 10 | 10 µS/cm | ± 10% | ± 10% | 80% |

* Note: Some test kits vary in sensitivity over the range of detection.  The specific range of readings is noted in parentheses.

Table 7.2. Data Quality Objectives for Nutrients Using Comparators

| Parameter | Method/range | Units | Detection Limit | Sensitivity* | Precision | Accuracy | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Nitrate Nitrogen | Zinc reduction | mg/l | 1.0 | 1.0 | ± 1.0 | ± 1.0 | 80% |
| Ortho-Phosphate | Ascorbic acid | mg/l | 0.2 | 0.2 (0-1.0) 0.5 (1.0-2.0) | ± 0.5 | ± 1.0 | 80% |

* Note: Some test kits vary in sensitivity over the range of detection.  The specific range of readings is noted in parentheses.

Table 7.3. Data Quality Objectives for Biological Parameters

| Parameter | Method/range | Units | Detection Limit | Sensitivity | Precision | Accuracy | Completeness |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Benthic Macro-invertebrates | Calif. Stream Bioassessment Protocol | N/A | Family level | N/A | < 10% difference | < 10% difference | 90% |
| Taxonomic ID’s | Calif. Stream Bioassessment Protocol | N/A | Family level | N/A | < 10% difference | < 10% difference | 90% |
| Physical Habitat | Calif. Stream Bioassessment Protocol | N/A | N/A | N/A | < 10% difference | < 10% difference | 90% |

NA:  not applicable

## 7.1.	Accuracy

7.1.1. Chemical and Physical Parameters

Accuracy describes how close the measurement is to its true value.  Accuracy is the measurement of a sample of known concentration and comparing the known value against the measured value.  The accuracy of chemical measurements will be checked by performing tests on standards each time equipment is checked out.  A standard is a known concentration of a certain solution.  Standards will be purchased from chemical or scientific supply companies or be prepared by a professional partner, e.g. a commercial or research laboratory. The concentration of the standards will be within the mid-range of the equipment.  The Data Quality Form: Accuracy, found in Appendix 1, will be used to record accuracy.

7.1.2. Biological Parameters

For benthic macroinvertebrate analysis, accuracy will be determined by having 10% of the samples (annually) re-analyzed and validated to CSBP Level 3 (10% to genus level) by a professional taxonomist.

7. 2.	Comparability

Comparability is the degree to which data can be compared directly to similar studies.  San Diego Stream Team will use the methods described in the following resource documents to ensure that their data can be compared to others:

- U.S. EPA’s Volunteer Monitoring Manuals for Streams, Lakes or Estuaries,

- SWRCB Clean Water Team Compendium for Water Quality Monitoring and Assessment, and

- California’s Department of Fish and Game’s (CDFG) California Stream Bioassessment Protocol (CSBP) for Citizen Monitors.

- Heal the Bay’s Malibu Creek Stream Team Pilot Project, Shattering the Myths of Volunteer Monitoring

- San Francisco Estuary Institute’s Volunteer Monitoring Protocols.

Before modifying these methods, or developing alternative or additional methods, technical advisors will evaluate and review the effects of the potential modification.  It will be important to address their concerns about data quality before proceeding with the monitoring program.

## 7. 3.	Completeness

Completeness is the fraction of planned data that must be collected in order to fulfill the statistical criteria of the project.  Volunteer data will not be used for legal or compliance uses.  There are no statistical criteria that require a certain percentage of data.  However, it is expected that 80% of all measurements could be taken when anticipated.  This accounts for adverse weather conditions, safety concerns, and equipment problems.

San Diego Stream Team will determine completeness by comparing the number of measurements that were scheduled to be collected to the number of measurements actually collected and deemed valid.  An invalid measurement is one that does not meet the sampling method requirements and/or the data quality objectives. Completeness results will be checked semiannually.  The Data Quality Form: Completeness, found in Appendix 1, will be used to record completeness.

## 7. 4.	Precision

7.4.1. Chemical and Physical Parameters

The precision objectives apply to duplicate and split samples taken as part of a Quality Control (QC) Session or as part of periodic in-field QC checks.  Precision describes how well repeated measurements agree.  The evaluation of precision described here relates to repeated measurements taken by either different volunteers on the same sample (at quality control sessions) or the same volunteer analyzing replicate samples (in the field).  Sampling variability will not be covered in this section.  The Data Quality Form: Precision, found in Appendix 1, will be used to record precision.

## 7. 5.	Representativeness

Representativeness describes how relevant the data is to the actual environmental conditions.  Problems can occur if:

- Samples are taken in a stream reach that does not describe the area of interest (e.g. a headwaters sample should not be taken downstream of a point source),

- Samples are taken in an unusual habitat type (e.g. a stagnant backwater instead of in the flowing portion of the creek),

- Samples are not analyzed or processed appropriately, causing conditions in the sample to change (e.g. water chemistry measurements are not taken immediately).

Representativeness will be ensured by processing the samples in accordance with Section 10, 11 and 12, by following the established methods, and by obtaining approval of this document.

## 7. 6.	Method Detection Limit and Sensitivity

The Method Detection Limit is the lowest possible concentration the instrument or equipment can detect.  This is important to record because we can never determine that a pollutant was not present, only that we could not detect it. Sensitivity is the ability of the instrument to detect one concentration from the next.  Detection Limits and Sensitivities are noted in Tables 7.1. - 7.3.

# 8.	Training Requirements

Field team leaders will receive a one-day training annually in physical and chemical monitoring field techniques.  Field team leaders will also participate in a semiannual training in benthic macroinvertebrate sampling (bioassessment).  The macroinvertebrate training will be provided by the California Department of Fish and Game, Regional Water Quality Control Board, Sustainable Lands Stewardship Institute, or other experts.  Trained field team leaders leaders may then train their rank-and-file citizen monitors. Individual trainees are evaluated by their performance of analytical and sampling techniques, by comparing their results to known values, and to results obtained by trainers and other trainees.

In addition to completion of the above described training courses, the field team leaders and citizen monitors must participate in semiannual Quality Control Sessions.  These Quality Control Sessions will be supervised by Quality Control Trainers and will provide an opportunity for participants to check the accuracy and precision of their equipment and techniques.  Quality Control Trainers are defined as water quality professionals from the U.S. Environmental Protection Agency, the State Water Resources Control Board, and the Regional Water Quality Control Boards.  Additional qualified trainers may be recruited and designated by the above agencies from experienced citizen monitoring organizations, universities and colleges, commercial analytical laboratories, and other federal, state, and local agencies.

Participants of the Quality Control Sessions will conduct duplicate tests on all analyses and meet the data quality objectives described in Section 7.  Monitors will also conduct tests on blind samples, where the concentration of the samples, known to the Quality Control Trainers, will be unknown to the monitors until measurements are determined. The Trainers will ensure that monitors are reading instruments and recording results correctly and that proper safety techniques are being used.  If a monitor does not meet the objectives, the trainers will re-train and re-test the monitor.  If there is insufficient time at the Quality Control Session to re-train and re-test monitors, the monitor will be scheduled for an additional training session.  The monitor will be encouraged to discontinue monitoring for the analysis of concern until training is completed.

Train-the-trainer workshops will be conducted by a professional taxonomist on an annual basis to certify new quality assurance leaders.   Quality assurance leaders will be responsible for overseeing quality control at the macroinvertebrate processing events.  A professional taxonomist will certify quality assurance leaders based on their ability to sort bethic macroinvertebrates from debris and their ability to identify aquatic insect larvae to family level.

Citizen monitors must attend semiannual trainings on the macroinvertebrate sorting and identification process.  Once the citizen monitors are trained in Standard Operating Procedure, it is the responsibility of the quality assurance leader to check their samples for 90% efficiency.

# 9.	Documentation and Records

All field results will be recorded at the time of completion, using the field data sheets (see Appendix 2).  Data sheets will be reviewed for outliers and omissions by the field team leader before leaving the sample site.   Completed data sheets will be given to the data manager and stored in hard copy form at the location specified in Section 5.2.  Field data sheets are archived for three years from the time they were collected.  If data entry is ever performed at another location, duplicate data sheets will be used, with the originals remaining at the headquarters site.  Hard copies of all data as well as computer back-up disks are maintained at headquarters.

All chain-of-custody forms, completed data quality control forms, permission slips from property owners, site location descriptions, and maintenance logs will also be kept at the headquarters location specified in Section 5.2. The maintenance log details the dates of equipment inspection, battery replacement and calibrations, as well as the dates that reagents and standards are replaced.

# 10.	Sampling Process Design

## 10.1.	Rationale for Selection of Sampling Sites

San Diego Stream Team generally follows the non-point pollution sampling design, as described in the manual “Measuring the Health of California Streams and Rivers: A methods manual for resource professionals, citizen monitors and natural resources students second edition” (Harrington and Born, 2000).

Sampling sites are indicated on the maps in Appendix 3.  The following criteria were evaluated when choosing sampling locations:

- access is safe,

- permission to cross private property is granted,

- sample can be taken in main river current or where homogeneous mixing of water occurs,

- sample is representative of the part of the water body of interest,

- location complements or supplements historical data,

- location represents an area that possesses unique value for fish and wildlife or recreational use.

- Site should be upstream of road, trail or other crossings if possible

- Site should consist of > 3 riffles per run-pool sequences if possible

Any reference sites are chosen upstream of any potential impact.  If a point pollution sampling design is required, a site chosen to reflect the impact of a particular discharge, tributary or land use is located downstream of the impact where the impact is completely integrated with the water, but upstream of any secondary discharge or disturbance.

Prior to final site selection, permission to access the stream was obtained from all property owners.  If access to the site becomes a problem, the field team leader will select a new site.  Safety issues are included in the Standard Operating Procedure.

Sample sites will be reviewed by the Field Coordinator and Technical Advisors before sending volunteers out to the site.  The field team leader will document permission and terms obtained from landowners, and will complete and file a Stream/Shore Walk form for the site, which will include a map and photographs.

## 10.2.	Sample Design Logistics

Volunteers are instructed to work in teams of at least two people.  If a scheduled team cannot conduct the sampling together, the field team leader is instructed to contact the Field Coordinator so that alternate arrangements can be made.

Prior to final site selection, permission to access the stream is obtained from all property owners.  If access to the site is a problem, the field team leader will select a new site following the site selection criteria identified in Section 10.1.

Safety measures will be discussed with all volunteers.  No instream sampling will be conducted if there are small creek flood warnings or advisories. It is the responsibility of the citizen monitoring organization to ensure the safety of their volunteer monitors.  Safety issues are included in the Standard Operating Procedure.

- Sampling Material Requirements

The manual “Measuring the Health of California Streams and Rivers: A methods manual for resource professionals, citizen monitors and natural resources students second edition” (Harrington and Born, 2000) describes the appropriate sampling procedure for collecting samples for water chemistry.  Water sampling apparatus may include hand held plastic containers.  Benthic invertebrates will be collected with a D shaped kick net (0.5 mm mesh) mounted on a pole.  Never use a sample bottle containing preservatives/fixing agents for sampling; in these cases always use a sampling device to collect the sample prior to transferring the sample into the bottle.

All samples are taken approximately in mid-stream, at least one inch below the surface.  If it is necessary to wade into the water, the sample collector stands downstream of the sample, taking a sample upstream.  If the collector disturbs sediment when wading, the collector will wait until the effect of disturbance is no longer present before taking the sample.

The following table describes the sampling equipment, sample holding container, sample preservation method and maximum holding time for each parameter.

Table 11.1 Sampling Material Requirements

| Parameter |  | Sample Bottle | Preferred / Maximum Holding Times |
| --- | --- | --- | --- |
| Conventional Parameters |  |  |  |
| Temperature |  | clear plastic bottle or sample directly | immediately |
| Dissolved oxygen |  | plastic bottle or sample directly | immediately |
| pH |  | plastic bottle or sample directly | immediately |
| Conductivity |  | plastic bottle or sample directly | immediately / refrigerate up to 24 hours |
| Nutrients |  |  |  |
| Nitrate N |  | plastic bottle or sample directly | immediately / refrigerate in dark for up to 48 hours |
| Phosphate |  | plastic bottle or sample directly | immediately |
| Biologicals |  |  |  |
| Benthic Macro-invertebrates | wide mouth plastic bottles |  | fix with ethanol immediately. Store indefinitely |

# 12.	Sample Handling and Custody Procedures

## 12.1.	Sample Handling

Identification information for each sample will be recorded on the field data sheets (see Appendix 2) when the sample is collected.   Samples that are not processed immediately in the field will be labeled with the waterbody name, sample location, sample number, date of collection and sampler’s name.

## 12.2.	Custody Procedures

The conventional water quality monitoring tests do not require specific custody procedures since they will, in most cases, be conducted immediately by the same person who performs the sampling.  In certain circumstances (such as driving rain or extreme cold), samples will be taken to a nearby residence for analysis.  Samples requiring chemical preservation, such as bioassessment samples, will be fixed prior to transport.

A Chain of Custody record must be maintained from the time the bioasessment sample is collected to its final disposition.  The record documents the transferring of samples from one volunteer to another member of the same organization, or from the citizen monitoring group to an outside professional laboratory.  Each transfer of custody must be noted and signed.  The individual responsible for the custody is to maintain direct control (e.g., possession or line of site) of the sample(s), or must maintain the sample(s) in a secured location, such as in a locked car.

The Chain of Custody record shall include (at a minimum) the following:

- Name of the waterbody,

- Sample location,

- Sample number,

- Sample date,

- Sample’s name and signature, and

- Preservative used (if any)

It also documents the date(s) and time(s) of transfer(s), and the name and signature of the sampler and the sample recipient.  When a professional lab performs quality control checks, their chain of custody forms and procedures are to be used.  In cases where the sample(s) remains in the custody of the monitoring organization, the field data sheet may be allowed to double as the chain of custody form.

## 12.3.	Disposal

All analyzed samples or spent chemicals (except for waste from the nitrate/cadmium reduction test) including used reagents, buffers or standards will be collected in a plastic bottle clearly marked “Waste” or “Poison”.  This waste material will be disposed of according to appropriate state and local regulations.  This will usually mean disposal into a drain connected to a sewage treatment plant.

Liquid waste from the cadmium reduction nitrate test will be kept separate and disposed of at a facility that is permitted to handle, transport, or dispose Cadmium (Cd) waste.

Waste disposal will be documented and documents retained in SDST files.

# 13.	Analytical Methods Requirements

Water chemistry is monitored using protocols outlined in the manual “Measuring the Health of California Streams and Rivers: A methods manual for resource professionals, citizen monitors and natural resources students second edition” (Harrington and Born, 2000).  The methods were chosen based on the following criteria:

- capability of volunteers to use methods,

- provide data of known quality,

- ease of use,

- methods can be compared to professional methods in AWWA Standard Methods.

If modifications of methods are needed, comparability will be determined by side-by-side comparisons with a US EPA or APHA Standard Method on no fewer than 50 samples.  If the results meet the same precision and accuracy requirements as the approved method, the new method will be accepted.

Table 13.1 outlines the methods to be used, any modifications to those methods, and the appropriate reference to a standard method.

Table 13.1  Analytical Methods for Water Quality Parameters

| Parameter | Method | Modification | Reference (a) |
| --- | --- | --- | --- |
| Temperature | Thermometric | Alcohol-filled thermometer marked in 0.5oC increments | 2550 B. |
| Dissolved Oxygen | Colorimetric indigo carmine Vacuum ampoules Color Comparator | none | ASTM D 888-87 |
| pH | Electrometric | Litmus indicator strips, non-bleeding | 4500-H B. |
| Conductivity | Electrometric | none | 2520 B. |
| Nitrate N | Cadmium Reduction or Zinc reduction Color Comparator | none | 4500 – NO3- E. |
| Benthic Macroinvertebrates | California Stream Bioassessment Protocol Level 2 (to Family) | none | Harrington, Jim, CDFG, 2000 |

All of the above methods, with the exception of dissolved oxygen via indigo carmine, pH via non-bleeding indicator strips, and benthic macroinvertebrates are described in Standard Methods for the Examination of Water and Wastewater 20th Edition.  American Public Health Association et al, 1998.

# 14.	Quality Control Requirements

Quality control samples will be taken to ensure valid data are collected.  Depending on the parameter, quality control samples will consist of blanks, replicate samples, and split samples.  In addition, quality control sessions (a.k.a. intercalibration exercises) will be held twice a year to verify the proper working order of equipment, refresh volunteers in monitoring techniques and determine whether the data quality objectives are being met.

## Cautions Regarding  Test Procedures

### 14.1.1.	Nutrients

The nitrate test measures nitrite as well as nitrate. Therefore the results for the nitrate test are actually mg/l Nitrite + Nitrate Nitrogen.  When mixing nitrate reagents take care not to agitate aggressively.  The LaMotte phosphate reagents have been shown to degrade well within their listed shelf life once opened.

## 14.2.	Replicates, Split Samples, and Standardization

Field Confirmations:  When a second method for measuring temperature, dissolved oxygen, and pH is available in the field, then the monitors are encouraged to perform both measurements on a split sample at least once daily.  Examples of this sort of redundant measurement would be:

- for temperature, the use of an electronic thermometer (such as those that are built into dissolved oxygen meters) and an armored thermometer;

- for dissolved oxygen, the use of an oxygen meter and an indigo carmine colorimetric kit;

- for pH, a meter and a non-bleeding indicator strip.

This will serve to provide backup capability if the more sensitive electronic meters fail, and will provide additional confidence as to the quality of the data. The results of both measurements will be recorded along with the procedure used on the field data sheet.  If both results are comparable then the result produced using the method of greater sensitivity will be the one entered in the final data set by the data manager in consultation with the field team leader.  If the two results are inconsistent, then the  field team leader will note on the data sheet which of the results will be entered on the final data set by the data manager.

Replicate Samples:  Replicate samples are two or more samples collected at the same time and place.  When there are only two replicates then these are referred to as duplicates.  For conventional water quality and nutrients duplicate field samples will be taken once every 20 samples or twice per sampling season, whichever comes first.  Duplicate samples will be collected as soon as possible after the initial sample has been collected, and will be subjected to identical handling and analysis. For benthic macroinvertebrate sampling, instead of duplicate sampling, each sampler will be evaluated annually by measuring the area sampled upstream of the net.  The area should be two square feet and should be verified by using a two square foot pvc frame.

Split Samples: Twice a Year, split spiked samples (standards) will be analyzed as part of the Quality Control Session.  The split standard is one sample, containing a known concentration of an analyte, that is divided equally into two or more sample containers.  Split standards will be analyzed by the volunteers, (except for dissolved oxygen, temperature, and pH), before the maximum sample handling time is exceeded.  Volunteers will analyze the split standard normally and will perform at least three analyses on that same sample. From these results accuracy and precision will be determined.

A minimum of 10% of the benthic macroinvertebrate samples will be subjected to taxonomic validation by an outside professional taxonomist. Following analysis by the citizen group the selected samples will be reconstituted and sent out for professional level 3 taxonomic analyses. Reconstituted means opening the vials containing the 100 identified specimens, pouring the specimens back into the original sample jar, and gently stirring the contents.  In addition, once a year citizen macroinvertebrate analysts will participate in an intercalibration exercise in which their subsampling/sorting and taxonomic skills will be evaluated.  A minimum of two teams of analysts will each inspect each other’s processed grids immediately following completion of the subsampling procedure.  There should be no more than 10% missed organisms.  A technical advisor should then evaluate each of the citizen analysts by testing their identification to order and family level on at least 20 specimens, including at least one representative from each of the major orders and families as determined by the technical advisor for that watershed. Accuracy and precision can be determined by the results of these validation and evaluation measures.

Standardization of Instruments and Procedures:  At the Quality Control Sessions the temperature measurements will be standardized by comparing our thermometers to a NIST-certified or calibrated thermometer in ice water and ambient temperature water.  All meters (pH, conductivity) will be evaluated at the Quality Control Session using standards provided with the assistance of a professional laboratory and/or the technical advisors.

Table 14.1 summarizes the quality control regimen.

Table 14.1  Summary of Quality Control Requirements

| Parameter | Blank | Duplicate Sample | Split Sample to lab | QC session |
| --- | --- | --- | --- | --- |
| Water quality |  |  |  |  |
| Temperature | none | 5% or a minimum of once a year | none | twice a year |
| Dissolved oxygen | none | 5% or a minimum of once a year | none | twice a year |
| pH | none | 5% or a minimum of once a year | none | twice a year |
| conductivity | daily | 5% or a minimum of once a year | none | twice a year |
| Phosphate | daily | 5% or a minimum of once a year | none | twice a year |
| Nitrate | daily | 5% or a minimum of once a year | none | twice a year |
| Biological Parameters |  |  |  |  |
| Benthic Invertebrates | none | None, instead conduct evaluation of sampling area annually | 10% per year | once a year |

# 15.	Instrument/Equipment Testing, Inspection and Maintenance

A maintenance log is kept by the  equipment manager.  This log details the dates of instrument and sampling gear inspection, calibrations performed in the laboratory, battery replacement, the dates reagents and standards are replaced, and any problems noted with instruments, samplers, or reagents. The equipment manager will examine kits for completeness of components: date, condition, and supply of reagents, and whether the equipment is in good working condition.

## 15.1.	Temperature

Before each use, thermometers are checked for breaks in the column.  If a break is observed, the alcohol thermometer will be placed in nearly boiling water so that the alcohol expands into the expansion chamber, and the alcohol forms a continuous column. verify accuracy by comparing with a calibrated or certified thermometer.

## 15.2.	Conductivity and pH

Before each use, conductivity and pH meters are checked to see if they are clean and in good working order.  Conductivity and pH meters are calibrated before each use.  Conductivity standards and pH buffers are replaced at least annually.  Conductivity standards are stored with the cap firmly in place and in a dry place kept away from extreme heat.  Do not re-use pH or conductivity standards.

## 15.3.	Nutrients

Before each use, test kits are checked to ensure that droppers, sample containers, and color comparators are clean and in working condition.  Reagents are replaced annually according to manufacturer’s instructions.

16.	Instrument Calibration / Standardization and Frequency

Instruments will be calibrated and reagents checked against standards according to the following schedule.  Standards will be purchased from a chemical supply company or prepared by (or with the assistance of) a professional laboratory.  Calibration records will be kept in the maintenance log at the headquarters location (described in Section 5.2.) where it can be easily accessed before and after equipment use.  Calibrations that are performed by monitors in the field are recorded on the field data sheets, also archived at the headquarters.  The frequency of calibration is described in Table 16.1.

Table 16.1  Instrument Calibration and Frequency

| Conventional Water Quality Parameters |  |  |
| --- | --- | --- |
| Equipment Type | Calibration Frequency | Standard or Calibration Instrument Used |
| Thermometric (Temperature) | Every 6 months | NIST calibrated or certified thermometer |
| pH meter | Every sampling day | pH 7.0 buffer and one other standard (4 or 10) |
| Conductivity meter | Every sampling day | Conductivity standard and distilled/deionized water |

| Nutrients (using comparators) |  |  |
| --- | --- | --- |
| Equipment type | Checked against Standard | Standard Used |
| Nitrate color comparator | every 6 months or when reagents replaced | nitrate standard |
| Phosphate color comparator | every 6 months or when reagents replaced | phosphate standards |

# 17.	Inspection/Acceptance Requirements

Upon receipt, buffer solutions, standards, and reagents used in the field kits will be inspected by the  equipment manager for leaks or broken seals, and to compare the age of each reagent to the manufacturer’s recommended shelf-life.  All other sampling equipment will be inspected for broken or missing parts, and will be tested to ensure proper operation.

Before usage, thermometers are inspected for breaks.  Breaks can be eliminated by heating (see Section 15.1).  If not, they will be returned to the manufacturer.

Reagents are replaced before they exceed manufacturer’s recommended shelf life.  These shelf lives are typically one to two years.  However, specific replacement dates can be determined by providing the reagent lot number to the manufacturer.  Reagent replacement dates are noted in the maintenance log.

# 18.	Data Acquisition Requirements

## 18.1.	Professional Analytical Data

Only certified analytical laboratories or academic laboratories (with approval of State and/or Regional Board staff) will be used for quality assurance checks and analysis of field samples.  The Technical Advisory Committee (TAC) or technical advisors will review these laboratories’ data as well as the volunteers.  They may also review the lab’s own quality control data to ensure data validity.

## 18.2.	Geographical Information/ Mapping

USGS maps will be used to verify watershed boundaries and river courses.  NOAA navigation charts can be used for mapping marine sampling sites.  Additional information on distribution of natural resources will be obtained from the National Park Service and the CDFG’s Biodiversity database.  Land use information will be obtained from local planning offices.  When information is requested, the agency will be asked to provide appropriate megadata and any information on data limitations.  This information will be maintained with the data files.

# 19.	Data Management

Field data sheets are checked and signed in the field by the  field team leader.  The  field team leader will identify any results where holding times have been exceeded, sample identification information is incorrect, samples were inappropriately handled, or calibration information is missing or inadequate.  Such data will be marked as unacceptable by the field team leader and will not be entered into the electronic data base by the data manager.

Independent laboratories will report their results to the  data manager.  The manager will verify sample identification information, review the chain-of-custody forms, and identify the data appropriately in the database.  These data are also reviewed by the technical advisors quarterly.

The data manager will review the field sheets and enter the data deemed acceptable by the  field team leader and the 
Technical Advisors.  Upon entering the data the data manager will sign and archive the field data sheets.  Data will be entered into a spreadsheet (MS Excel) or a database (MS Access) in a way that will be compatible with EPA’s STORET and the Regional WQCB’s database guidelines.  Following initial data entry the data coordinator will review electronic data, compare to the original data sheets and correct entry errors.  After performing data checks, and ensuring that data quality objectives have been met, data analysis will be performed.

Raw data will be provided to the State WRCB and Regional WQCB in electronic form at least once every two years so that it can be included in the 305(b) report.  Appropriate quality assurance information may be provided upon request.

# 20.	Assessment and Response Actions

Review of all field and laboratory data activities is the responsibility of the  data manager, with the assistance of the technical advisory committee.  Volunteers will be accompanied by  a field team leader, or a technical advisor on all sampling trips .  If possible, volunteers in need of performance improvement will be retrained on-site.  All volunteers must attend a refresher course offered by the citizen monitoring group.  If errors in sampling technique are consistently identified, retraining may be scheduled more frequently.

Within the first three months of the monitoring project, the State Water Board or Regional Board staff, or its designee, will evaluate field and laboratory performance and provide a report to the citizen monitoring group.  All field and laboratory activities, and records may be reviewed by State and EPA quality assurance officers as requested.

# 21.	Reports

The technical advisors will review draft reports to ensure the accuracy of data analysis and data interpretation.  Raw data will be made available to data users per their request.  The citizen monitoring organization(s) will report their data to its (their) constituents after quality assurance has been reviewed and approved by their technical advisors.  Every effort will be made to submit data and/or a report to the State and/or Regional Board staff in a fashion timely for their data uses, e.g. 305(b) reports.

# 22.	Data Review, Validation and Verification

Data sheets or data files are reviewed quarterly by the technical advisors to determine if the data meet the Quality Assurance Project Plan objectives.  They will identify outliers, spurious results or omissions to the  data manager.  They will also evaluate compliance with the data quality objectives.  They will suggest corrective action to the data manager that will be implemented by the  field and laboratory coordinators.  Problems with data quality and corrective action will be reported in final reports.

# 23.	Validation and Verification Methods

As part of standard field protocols, any sample readings out of the expected range will be reported to the  data manager.  A second sample will be taken as soon as possible to verify the condition.  If the data is invalid, then the data will be noted (flagged) on the data sheet.  We will take further actions to trace the sources of error, and to correct those problems.   If the error is a result of improper monitoring procedures, then we may re-train monitors until their performance is acceptable.

It is the responsibility of the citizen monitoring group to re-train volunteers until performance is acceptable.

# 24.	Reconciliation with DQOs

The Technical Advisory Committee working with the  data manager will review data quarterly to determine if the data quality objectives (DQOs) have been met.  A quorum of 1/2+1of the technical advisory committee will be required for committee decisions.  If a quorum is not met at the meeting, work will still proceed. The work product (e.g., review and comments on data or reports) will then be sent out to the whole technical advisory committee for approval with a 30-day review period.

If data do not meet the project’s specifications, the following actions will be taken.  First, the technical advisors working with the  data and equipment managers will review the errors and determine if the problem is equipment failure, calibration/maintenance techniques, or monitoring/sampling techniques. They will suggest corrective action.  If the problem cannot be corrected by training, revision of techniques, or replacement of supplies/equipment, then the technical advisors and the TAC will review the DQOs and determine if the DQOs are feasible. If the specific DQOs are not achievable, they will determine whether the specific DQO can be relaxed, or if the parameter should be eliminated from the monitoring program.  Any revisions to DQOs will be appended to this QA plan with the revision date and the reason for modification.  The appended QAPP will be sent to the quality assurance panel that approved and signed this plan.  When the appended QAPP is approved, all data meeting the new DQOs will be entered into the database.  Archived data can also be entered.

APPENDIX 1.  Data Quality Forms

Use these forms to keep records of your own quality assurance.  There are data quality forms for accuracy, completeness and precision.  Temperature, Dissolved Oxygen, pH, and Conductivity are already listed under Parameters/units.  Additional space is available for other parameters.

# Data Quality Form: Accuracy	 	Quality Control Session

| San Diego Stream Team | Type of Session (field or lab) |
| --- | --- |
| Your Name | Quality Assurance Leader |
| Date |  |

| Parameter/ units | Sensitivity | Accuracy Objective | Standard Conc. | Analytical Result | Estimated Bias | Meet Objective? Yes or No | Corrective action planned | Date Corrective Action taken |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature O C |  |  |  |  |  |  |  |  |
| Dissolved Oxygen (mg/l) |  |  |  |  |  |  |  |  |
| PH Standard units |  |  |  |  |  |  |  |  |
| Conductivity (umhos/cm) |  |  |  |  |  |  |  |  |

Comments:

# Data Quality Form: Completeness	 	Quality Control Session

| San Diego Stream Team |  |  | Type of Session (field or lab) |  |  |
| --- | --- | --- | --- | --- | --- |
| Your Name |  |  | Quality Assurance Leader |  |  |
| Date |  |  |  |  |  |
| Parameter | Collection Period | No. of Samples Anticipated |  | No. Valid Samples Collected and Analyzed | Percent Complete |
| Temperature o C |  |  |  |  |  |
| Dissolved Oxygen (mg/l) |  |  |  |  |  |
| pH standard units |  |  |  |  |  |
| Conductivity (umhos/cm) |  |  |  |  |  |

Comments:

# Data Quality Form: Precision	 	Quality Control Session

| San Diego Stream Team | Type of Session (field or lab) |
| --- | --- |
| Your Name | Quality Assurance Leader |
| Date |  |

| Parameter/ units | Mean (x) | Standard Deviation (s.d.) | s.d./x | Precision Objective | Meet Objective? Yes or No | Corrective action planned | Date Corrective Action taken |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature o C |  |  |  |  |  |  |  |
| Dissolved Oxygen mg/l |  |  |  |  |  |  |  |
| pH standard units |  |  |  |  |  |  |  |
| Conductivity (umhos/cm) |  |  |  |  |  |  |  |

Comments:

APPENDIX 2.  Data and Observation Sheets

Design and insert your own forms here, based on the information discussed in your QAPP.

APPENDIX 3.  Maps of Sampling Sites

Insert maps of your sampling sites here.