## 2003 DATA QAULITY REPORT

Data quality is defined by a series of statements that set standards the data must meet. These standards cover the way the sample is collected and analyzed. In general, certain performance criteria (data quality objectives) are established to ensure that the data are acceptable and usable. Three quantitative performance criteria were used in the Program to evaluate the degree of uncertainty or usability of the data to users. These criteria included precision, accuracy, and completeness.  Precision is expressed as relative percent difference (%RPD) between two duplicate samples. Accuracy is expressed as the percent recovery (REC%) of an analyte spike. Completeness is defined as the percentage of actual measurements, which are judged to be valid, over the planned overall measurements.

The following provides the discussion of analytical laboratory QA/QC results.

The data used were obtained from measurements of laboratory duplicates or spiked samples. The number of data sets ranged from 4 to 31 with various concentrations. Although the precision (%RPD) for oil and grease was slightly higher than the DQO (Table 1), it was

Table 1. Data Quality Objectives (DQO) and levels achieved for analytical parameters

| Parameter | Precision (RPD%) |  | Accuracy (REC%) |  | Completeness (%) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | DQO | Achieved | DQO | Achieved | DQO | Achieved |
| Diazinon | <25 | 0-15 | 60-117 | 76-103 | >95 | 100 |
| Chlorpyrifos | <25 | 0-23 | 54-125 | 72-110 | >95 | 100 |
| Dissolved Cu | <20 | 0-16 | 75-125 | 85-102 | >95 | 100 |
| Dissolved Cd | <20 | 0-6 | 75-125 | 97-112 | >95 | 100 |
| Dissolved Pb | <20 | 0-8 | 75-125 | 80-130 | >95 | 100 |
| Dissolved Zn | <20 | 0.9-19 | 75-125 | 92-113 | >95 | 100 |
| Hardness | <20 | 0-8 | 75-125 | ND | >95 | 100 |
| Oil & Grease | <20 | 1-22 | 80-120 | 80-110 | >95 | 100 |
| MBAS | <20 | 0-11 | 80-120 | 90-120 | >95 | 100 |
| Ammonia | <20 | 1-13 | 80-120 | 67-101 | >95 | 100 |
| Nitrate | <20 | 0-14 | 80-120 | 14-196 | >95 | 100 |
| Nitrite | <20 | 0-12 | 80-120 | 80-120 | >95 | 100 |
| TKN | <20 | 0-7 | 80-120 | 63-112 | >95 | 100 |
| Orthophosphate | <10 | 0-11 | 80-120 | 90-120 | >95 | 100 |
| Total P | <10 | 0-11 | 80-120 | 92-116 | >95 | 100 |
| Total coliform | <50 | 0-46 | N/A | ND | >95 | 98 |
| Fecal coliform | <50 | 0-46 | N/A | ND | >95 | 98 |
| Enterococcus | <50 | 0-32 | N/A | ND | >95 | 98 |

acceptable as the QC result was flagged as QR-04 (Table 2). Accuracies for dissolved lead, ammonia, nitrate, and total Kjeldahl nitrogen (TKN) were lower than DQOs, but they were acceptable as the QC results were flagged as QM-4X or QM-05 (Table 2).  A few of the 95 water samples analyzed for total phosphorus (TP) had lower results than those for orthophosphate, which was likely due to leaking test-tube caps in the laboratory.  After replacing those faulty caps, samples were reanalyzed for total phosphorus, and the reanalysis results were higher and reported.

The laboratory analysis for nitrate generated low results when the nitrate concentration was greater then 5 mg/L NO3 as N. It was determined that the cadmium column was unable to reduce all the nitrates when the nitrate concentration was greater than 5 mg/L NO3 as N. The County and contract laboratory worked closely together to improve the accuracy of measuring nitrate concentrations in water samples. Further discussion can be found in the subsequent section.

Table 1. Data qualifiers and descriptions used by the contract laboratory

| Qualifier | Description |
| --- | --- |
| QM-4X | The spike recovery was outside of the QC acceptance limits for the MS and/or MSD due to analyte concentration at 4 times or greater the spike concentration. The QC batch was accepted based on LCS and/or LCSD recoveries within the acceptance limits. |
| QR-04 | The RPD between the sample and sample duplicate is not valid since both results are below the reporting limit for this analyte. |
| HT-03 | This analysis was originally performed within the recommended EPA holding time. Data review determined that reanalysis was necessary. |
| QM-05 | The spike recovery was outside acceptance limits for the MS and/or MSD due to matrix interference. The LCS and/or LCSD were within acceptance limits showing that the laboratory is in control and the data is acceptable. |
| W-02 | The sample for nitrate analysis was preserved with H2SO4 after the nitrite portion of the analysis was completed to extend the holding time for the sample. Nitrate results are corrected for the nitrite contribution per the method. |

Two out of 128 water samples for bacterial indicators were received by the laboratory just after the 6-hour holding time. The data were flagged internally.

### Laboratory Method Improvement for Nitrate Analysis

Results from the field measurements of nitrate were not comparable with those from the laboratory analysis, particularly when field measured nitrate concentrations were greater than 5 mg/L NO3 as N (Figure 1). These results suggested a problem with methodology. Initially the accuracy of the field test kit was examined and determined to be meeting the DQOs. Examination into the accuracy of the laboratory method determined that the laboratory method was not meeting the DQOs. A description of this investigation and actions taken are described below.

Figure 1. Comparison of nitrate results from field and laboratory measurements

#### Examination of the Field Method

In an effort to determine if a methodology problem was occurring in the field test method due to matrix interference or other factors, a series of nitrate samples were prepared by spiking a water sample collected in the field with a stock nitrate standard solution. These samples were analyzed for nitrate using, Nitrate 3 test kit (CHEMetrics, Inc., Calverton, VA) used by the County dry weather monitoring program, which has a method detection limit of 6.0 mg/L NO3 with a measurement range of 0-70 mg/L NO3.

Nitrate results for the spiked samples using the test kit method indicated the relative percent difference (%RPD) of <10% compared to 13% RPD for the non-spiked sample. The percent recovery of the spike ranged from 112% to 116%, which is within the goals of the DQO. The results demonstrated that the test kit method provided data with expected precision and accuracy.

#### Examination of the Laboratory Method

In an investigation of the analytical method, County staff met with the contract laboratory’s QA director to discuss the discrepancy in nitrate results between the test kit and laboratory methods.  The contract laboratory agreed to perform method verification analyses to evaluate the laboratory nitrate method.

Upon reviewing the analyst’s procedure, it was determined that the cadmium reduction column used by the laboratory might not effectively reduce the entire amount of nitrate in samples when nitrate concentrations are 5 mg/L as N or greater. The inefficiency of this reduction might have resulted in low concentrations when compared to those determined by the test kit method. It was determined that samples above this concentration require a dilution step prior to passing through the cadmium reduction column.

Laboratory staff analyzed the spiked samples prepared and submitted by County staff. The analysis was performed using both the old standard operating procedure (SOP) and the new SOP to verify that the problem had been resolved. The following table summarizes the nitrate results (reported as N) for the spiked samples. As expected, the old SOP produced nitrate results with a low recovery for samples greater than 5 mg/L as N. The new SOP greatly improved the accuracy, but there were still two of the spiked samples below the DQO (Table 3).

Table 3. Nitrate results for spiked samples analyzed by the Lab

| Spiked Sample | Old SOP Results | Old SOP Accuracy | New SOP Results | New SOP % Accuracy |
| --- | --- | --- | --- | --- |
| (mg/L) | (mg/L) | (REC%) | (mg/L) | (REC%) |
| 1.81 | 1.62 | 89 | ND | ND |
| 6.72 | 4.75 | 71 | 5.45 | 81 |
| 11.63 | 8.24 | 71 | 7.45 | 64 |
| 16.53 | 10.54 | 64 | 14.4 | 87 |
| 21.44 | 7.14 | 33 | 15.8 | 74 |
| 31.26 | 6.18 | 20 | 30.6 | 98 |

In reviewing the process with the laboratory staff, it became evident that additional dilutions were required to improve the accuracy. New spiked samples were prepared to validate the new SOP (concentrations of 10, 15, and 30 mg/L as N). The results of this validation using the new SOP and additional dilutions resulted in a percent recovery ranging from 92 to 96%. As a result, the Lab revised its standard operating procedure to incorporate a sample screening evaluation and the requirement of performing sample dilutions prior to passing samples through the cadmium reduction column (EMA, 2003).

The resultant action from these methods examination was that the laboratory agreed to analyze water samples with inaccurate results for nitrate determined previously using the old SOP. A total of 24 water samples were recollected by staff for nitrate analysis using the new SOP.  A plot of the samples analyzed after the SOP resulted in a linear correlation with a slope of 0.95 and r2 of 0.89.